Hudson River (HUD) NERR Meteorological Metadata

January 2002 - December 2002

Latest Update: August 23, 2024

I. Data Set & Research Descriptors

1.   Principal Investigator: William C. Nieder, Research Coordinator

Contact Persons:

William C. Nieder, Research Coordinator, email: wcnieder@gw.dec.state.ny.us

Erik Bedan, Research Assistant, email: esbedan@gw.dec.state.ny.us

Address:

Hudson River NERR

c/o Bard College Field Station

Annandale, NY 12504

Phone: 845-758-7010

Fax: 845-758-7033

2. Entry verification:

Person responsible for data verification: Erik Bedan, Research Assistant

a)	 Data Input Procedures:

The 15-minute, 1-hour average, and 24-hour data were downloaded from each

instrument on the weather station to a Campbell Scientific CR10X datalogger.

The CDMO Datalogger Program (hud30.csi) was loaded into the CR10X, which

controls the sensors and the data collection schedule (see 2b of the Entry

Verification section for the data collection schedule).  For data storage, the

CR10X stored the data, via cable, on a Campbell Scientific SM 192 storage

module.  The storage module was manually retrieved at the end of every month.

Using the PC208W software supplied by Campbell Scientific, the data were

uploaded from the storage module and stored on computer.  All raw data files

were archived at the CDMO and at the reserve on ditto tapes, and CD-ROM.

Anomalous or missing data were either deleted or identified in sections 11 & 12.

After the data upload, the CDMO Weather Data Management Program

(WDMP) was used to convert the data into an Access database.  This program was

developed in Visual Basic to interface with the NERRS data collection schedule

(see 2b of the Entry Verification section for the data collection schedule).

The WDMP will automatically input and convert the monthly raw data file into an

Access Database.  There were three main tasks completed by the WDMP.  First, it

converted the comma delimited monthly raw data file into an Access Database.

Secondly, it checked the data against a predetermined set of error criteria (see

Part C of this section).  Finally, it produced error and summary reports.  Any

anomalous data were investigated and were noted below in Anomalous Data section.

Any data corrections that were performed were noted in the Data Correction

section below.  Most common errors from monthly reports were wind speeds below

the 0.5m/s criteria.

i) Data is collected in the following formats:

1) 15 minute data are instantaneous readings except for PAR and

precipitation data that are totalized from 5 second samples sorted by date and

time. 	(Arrays 150 and 151)

2) Hourly averages (Arrays 101 and 102) are calculated from 5 second

samples sorted by date and time except for PAR and precipitation

data that are hourly totals calculated from 15 minute totals (Arrays 105 and

106).

3) Daily average (arrays 241 and 242), maximum with time, and

minimum with time (arrays 243 and 244) are calculated from 5 second samples

sorted by date and time except for PAR and precipitation data which are 24

hour totals calculated from hourly totals (arrays 245 and 246).

ii) 15 minute sample point parameters: Date, Time, Air Temperature (°C),

Relative Humidity (%), LiCor (PAR), Barometric Pressure (mb), Wind Speed (m/s),

Wind Direction (Array 150); Rainfall (mm) (Array 151)

iii) Hourly average parameters: Date, Time, Air Temperature (°C), Relative

Humidity (%), Barometric Pressure (mb) (Array 101); Wind Speed (m/s), Wind

Direction, Wind Speed	Maximum (Array 102)

iv) Hourly total parameters: LiCor (PAR) (Array 105); Rainfall (mm) (Array

106)

v) Daily Average parameters: Date, Time, Air Temperature (°C), Relative

Humidity (%), Barometric Pressure (mb) (Array 241); Wind Speed (m/s), Wind

Direction, Wind Direction Standard Deviation (using Yamartino's Algorithm) (Array

242)

vi) Daily Total parameter: LiCor (PAR) (Array 245); Rainfall (mm) (Array

246)

vii) Daily Maximum parameters: Date, Time, Air Temperature (°C), Time,

Relative Humidity (%), Time, LiCor (PAR), Time, Barometric Pressure (mb), Time,

Wind Speed (m/s), Time, Battery Voltage, Time (Array 243)

viii) Daily Minimum parameters: Date, Time, Air Temperature (°C), Time,

Relative Humidity (%), Time, LiCor (PAR), Time, Barometric Pressure (mb), Time,

Wind Speed (m/s), Time, Battery Voltage, Time (Array 244)

c) Error/Anomalous Data Criteria

Air Temp:

- 15 min sample greater than max for the day

- 15 min sample less than the min for the day

- 15 min sample greater than 3.0 °C from the previous 15 minutes

- Max and Min values not recorded for the day

- 1-hour average greater than 10% above the greatest 15 min sample

recorded in the hour

Relative Humidity:

- Changed by more than 25% from the previous 15 minutes

- Max and Min values not recorded for the day

- 1-hour average greater than 10% above the greatest 15 min sample

recorded in the hour

Rainfall:

- Precipitation greater than 5 mm in 15 minutes

- No precipitation for the month

Wind Speed:

- Wind speed greater than 30 m/s

- Wind speed less than 0.5 m/s

Wind Direction:

- Wind direction greater than 360 degrees

- Wind direction less than 0 degrees

Pressure:

- Pressure greater than 1040 mb or less than 980 mb

- Pressure changes greater than 5 mb per hour

- Max and Min values not recorded for the day

- 1-hour average greater than 10% above the greatest 15 min sample

recorded in the hour

Time:

- 15-minute interval not recorded

For all data:

- Duplicate interval data

3. Research objectives:

The objective of this study is to monitor the meteorological conditions at

the Tivoli Bays component site of the Hudson River National Estuarine Research

Reserve.  Measurements are taken all year long at one location in accordance

with section 2b of the Entry Verification section.  A water quality-monitoring

program has been ongoing for ten years at this component site, and the

meteorological data will help provide ancillary data to that program.  In turn,

this will help us to better understand the relationships between the atmospheric

and aquatic environments at this site.

4. Research methods:

For data collection, the CR10X datalogger is programmed at intervals that

are in accordance with section 2b of the Entry Verification section.  Data are

retrieved and stored from the datalogger via the Campbell Scientific storage

module at the end of each month.  The data are analyzed using only the Weather

Data Management Program (WDMP).  The program creates a database and error report

for each month of raw data.   The error reports are used to aid in the QA/QC of

the data, along with the guidelines in section 2c.  Any errors are reported and

corrected (if necessary) in sections 11 & 12.	For routine maintenance,

sensors are investigated at least once a week to ensure there is no damage or

blockage to the sensors.  According to Campbell Scientific, for proper data

collection, sensors are to be calibrated every two 2 years.  After that time,

the sensors are removed from the tower, and shipped to their respective

manufacturer.  The calibration methods are performed by the manufacturers for

proper and professional calibration.  In order to avoid missing data, a second

set of sensors will be calibrated and available to be installed on the tower to

continue data collection.

5. Site location and character:

The Hudson River National Estuarine Research Reserve (HUDNERR) is a multi-

component site composed of approximately 5,000 acres of tidal wetland, open

water, and adjacent upland buffer habitat.  The four component sites of the

Reserve are located on the Hudson River in New York State by river miles (RM)

proceeding north from the southern tip of Manhattan.  These include: Stockport

Flats (RM 124)(42°02'30"N 73°46'00"W), Columbia County; Tivoli Bays (RM

98)(42°02'15"N 73°55'10"W), Dutchess County; and Iona Island (RM 45)(41°18'15"N

73°58'45"W) and Piermont Marsh (RM 24)(41°02'30"N 73°54'15"W),Rockland County.

These four components are located along 100 river miles of the Hudson River and

are representative of the diverse plant and animal communities that occupy the

salinity gradient within the Hudson estuary.  Development within the watersheds

of Hudson River NERR component sites ranges from predominantly urban/suburban to

forested/agricultural.

The weather station (FS) is located at the Tivoli Bays component site in

Annandale, NY (42°01'05.46"N 73°55'01.13"W).  A 30 foot, aluminum tower is used

to elevate some of the weather monitoring equipment. The tower is on the deck of

the office building, on the west side, 9 feet off the ground.  The wind

anemometer, wind speed, and light sensors are all located at the top of the

tower, 39 feet off the ground.  Both the datalogger and barometric pressure

sensor are enclosed within a fiberglass case attached to the tower, 12 feet off

the ground.  A heated rain gauge is next to the tower, attached to the building,

16 feet off the ground.  The temperature/humidity sensor is next to the tower,

attached to the deck handrail, 12 feet off the ground.  Although trees surround

the area, the tree line begins approximately 60 feet from the tower in most

directions.  The trees are at similar heights to the tower, but the sensors are

not shaded at that location.  The tower is approximately 1.2 miles Southeast of

the TS water quality monitoring station, 2.3 miles Southeast of the TN water

quality monitoring station, and 0.2 miles Northwest of the SK water quality

monitoring station.

6. Data collection period:

The first data were recorded on January 14, 2002 at 16:45 hours EST, and ended

on December 31, 2002 at 2400 hours EST.

7. Distribution:

According to the Ocean and Coastal Resource Management Data Dissemination

Policy for the NERRS System-wide Monitoring Program, NOAA/ERD retains the right

to analyze, synthesize and publish summaries of the NERRS System-wide Monitoring

Program data.  The PI retains the right to be fully credited for having

collected and processed the data.  Following academic courtesy standards, the PI

and NERR site where the data were collected will be contacted and fully

acknowledged in any subsequent publications in which any part of the data are

used.  Manuscripts resulting from the NOAA/OCRM supported research that are

produced for publication in open literature, including refereed scientific

journals, will acknowledge that the research was conducted under an award from

the Estuarine Reserves Division, Office of Ocean and Coastal Resource

Management, National Ocean Service, National Oceanic and Atmospheric

Administration.  The data set enclosed within this package/transmission is only

as good as the quality assurance/quality control procedures outlined by the

enclosed metadata reporting statement.  The user bears all responsibility for

its subsequent use/misuse in any further analyses or comparisons.  The Federal

government does not assume liability to the Recipient or third persons, nor will

the Federal government reimburse or indemnify the Recipient for its liability

due to any losses resulting in any way from the use of this data.

NERR weather data and metadata can be obtained from the Research

Coordinator at the individual NERR site (please see Section 1 Principal

nvestigators and contact persons), from the Data Manager at the Centralized Data

Management Office (please see personnel directory under the general information

link on the CDMO home page) and online at the CDMO home page

http://cdmo.baruch.sc.edu.  Data are available in text format and

Access data tables.

8. Associated researchers and projects:

The Hudson River NERR water quality-monitoring program examines the

physical and chemical constituents of the tributary waters and the tidal waters

discharging into the HUDNERR marshes and marsh waters.  Measurements include

seston, dissolved oxygen (DO), alkalinity, pH, temperature, salinity, and

conductivity. In addition, samples are analyzed for nutrient concentrations of

nitrate, phosphate, sulphate, and chloride.

Associated researchers with work at Tivoli Bays include: scientists from

the Institute of Ecosystem Studies, Millbrook, NY; Yale School of Forestry and

Environmental Studies, New Haven, CT; and Rensselaer Polytechnic Institute,

Troy, NY.

II. Physical Structure Descriptors

9. Sensor specifications, operating range, accuracy, date of last calibration

Sensors in use from July 2001-June 2002:

Li-Cor Quantum Sensor

Model # LI190SB

Stability: <±2% change over 1 yr

Operating Temperature: -40 to 65°C

Sensitivity: typically 5 µA per 1000µmoles s-1 m-2

Light spectrum wavelength: 400-700 nm

Date of last calibration: December 2000

Wind Sentry

Model # 03001

Operating Range: 0-50 m/s; 360° mechanical

Date of last calibration: December 2000

Temperature and Relative Humidity

Model #:  HMP35C

Operating Range: -35 to 50°C

Temperature Accuracy: ± 0.2 °C @ 20°C

Relative Humidity Measurement Range: 0-100% non-condensing

RH Accuracy: ±2% RH (0-90%) and ±3%(90-100%)

Uncertainty of calibration: ±1.2% RH

Date of Last calibration: December 2000

Barometric Pressure Sensor

Model # CS105

Operating Range: 600 to 1060 mb

Operating Temperature: -40 to 60°C

Humidity: non-condensing

Accuracy: ±0.5 to 6.0 mb (+20 to 60°C)

Stability: ± 0.1 mb per year

Date of Last calibration: December 2000

Tipping Bucket Rain and Snow Gauge

Model #:  375-L

Operating Range: -20°C to 50°C; 0-100% RH

Sensitivity: 0.1mm

Accuracy: ±0.5% at <1.25/hr

Date of Last calibration: December 2000

The following sensors in use from July 2002-February 2003:

Li-Cor Quantum Sensor

Model # LI190SB

Stability: <±2% change over 1 yr

Operating Temperature: -40 to 65°C

Sensitivity: typically 5 mV per 1000 micromoles s-1 m-2

Light spectrum wavelength: 400-700 nm

Date of last calibration: May 2002

Wind Sentry

Model # 03001

Operating Range: 0-50 m/s; 360° mechanical

Date of last calibration: May 2002

Temperature and Relative Humidity

Model #:  HMP45C

Operating Range: -35°C to 50°C

Temperature Accuracy: ± 0.2 °C @ 20°C

Relative Humidity Measurement Range: 0-100% non-condensing

RH Accuracy: ±2% RH (0-90%) and ±3%(90-100%)

Uncertainty of calibration: ±1.2% RH

Date of Last calibration: May 2002

Barometric Pressure Sensor

Model # CS105

Operating Range: 600-1060 mb

Operating Temperature: -40°C to 60°C

Humidity: non-condensing

Accuracy: ±0.5 to 6.0 mb (+20°C to 60°C)

Stability: ± 0.1 mb per year

Date of Last calibration: May 2002

Tipping Bucket Rain and Snow Gauge

Model #: 375-L

Operating Range: -20°C to 50°C; 0-100% RH

Sensitivity: 0.1mm

Accuracy: ±0.5% at <1.25/hr

Date of Last calibration: May 2002

10. Coded variable indicator and variable code definitions:

FS - Field Station at Tivoli Bays

11. Data anomalies:

Arrays:

During 2022 all pre-2007 weather data were revisited by the CDMO. Historically those datasets included 15 minute, hourly (60), and daily data arrays (144). As directed by the NERRS Data Management Committee, the CDMO removed the hourly and daily data arrays leaving only the 15 minute data to make the entire NERRS SWMP weather dataset consistent in its reporting. All references to the 60 and 144 arrays were left in the metadata document as they may still provide valuable information, but users should be aware that they are largely no longer relevant. The updated datasets were uploaded to the database and made available through the various data applications at www.nerrsdata.org/get/landing.cfm throughout the fall of 2022.

PAR flagging update August 2024:

During SWMP data synthesis work it was noted that PAR values were anomalously high beginning in 2001 and returned to more expected values following the sensor swap on June 3, 2002. PAR data flagging was changed by the CDMO to suspect from January 10, 2001 15:15 through June 3, 2002 07:45.

January 2002

At this site, wintery weather requires us to utilize a heating tipping bucket

for precipitation data. The heater may not have been functioning this month,

precipitation data are suspect.

Array 	Day	Julian	Time	Error

102	14	14	1700	Wind speed is less than 0.5 m/s from  14 ( 14)

1700 to  15 ( 15)  800

102	19	19	1700	Wind speed is less than 0.5 m/s from  19 ( 19)

1700 to  20 ( 20)  900

102	30	30	2300	Wind speed is less than 0.5 m/s from  30 ( 30)

2300 to  31 ( 31)  2400

February 2002

The following error occured because the "1 hour data" is averaged from 5-second

readings taken during that hour while the "15 minute data" (in this case the 15

minute maximum) is not averaged, and only recorded at 15-minute intervals.  The

data were left in.

Array 	Julian	Day	Time	Error

101	44	13	1500	Air temp average in 1 hour data ( .02) is greater than

15 minute maximum (-.00712)

Array 	Day	Julian	Time	Error

102	1	32	100	Wind speed is less than 0.5 m/s from  1 ( 32)

100 to  1 ( 32)  1300

102	6	37	2200	Wind speed is less than 0.5 m/s from  6 ( 37)

2200 to  7 ( 38)  1000

102	11	42	1900	Wind speed is less than 0.5 m/s from  11 ( 42)

1900 to  12 ( 43)  800

102	16	47	1600	Wind speed is less than 0.5 m/s from  16 ( 47)

1600 to  17 ( 48)  700

March 2002

Data points were reviewed before and after the following errors. There was a

notable change, but the data were consistent.  It was most likely an event.  The

data were left in.

Array 	Julian	Day	Time	Error

150	62	3	1800	Air temp difference from  3 ( 62) 1800 ( 13.667) to  3 (

62) 1815 ( 10.559) is greater than 3.0 degrees C

150	68	9	1315	Air temp difference from  9 ( 68) 1315 ( 14.308) to  9 (

68) 1330 ( 17.684) is greater than 3.0 degrees C

150	80	21	1915	Air temp difference from  21 ( 80) 1915 ( 9.735) to  21

( 80) 1930 ( 5.9559) is greater than 3.0 degrees C

Data points were reviewed before and after the following error. There was a

notable change, but the data were consistent.  It was most likely an event.  The

data were left in.

Array 	Julian	Day	Time	Error

150	62	3	2030	Rel hum difference from  3 ( 62) 2030 ( 46.337) to  3 (

62) 2045 ( 73.16) is greater than 25%

Array 	Day	Julian	Time	Error

102	12	71	1800	Wind speed is less than 0.5 m/s from  12 ( 71)  1800

to  13 ( 72)  1000

102	13	72	1600	Wind speed is less than 0.5 m/s from  13 ( 72)  1600

to  14 ( 73)  900

102	14	73	1800	Wind speed is less than 0.5 m/s from  14 ( 73)  1800

to  15 ( 74)  600

102	17	76	1700	Wind speed is less than 0.5 m/s from  17 ( 76)  1700

to  18 ( 77)  1500

102	18	77	1700	Wind speed is less than 0.5 m/s from  18 ( 77)  1700

to  19 ( 78)  800

102	19	78	1700	Wind speed is less than 0.5 m/s from  19 ( 78)  1700

to  21 ( 80)  1100

April 2002

Data points were reviewed before and after the following errors. There was a

notable change, but the data were consistent.  It was most likely an event.  The

data were left in.

Array 	Julian	Day	Time	Error

150	107	17	1015	Air temp difference from  17 ( 107) 1015 ( 26.043) to

17 ( 107) 1030 ( 29.052) is greater than 3.0 degrees C

150	109	19	1200	Air temp difference from  19 ( 109) 1200 ( 29.947) to

19 ( 109) 1215 ( 26.558) is greater than 3.0 degrees C

150	109	19	1215	Air temp difference from  19 ( 109) 1215 ( 26.558) to

19 ( 109) 1230 ( 19.046) is greater than 3.0 degrees C

Data points were reviewed before and after the following error. There was a

notable change, and some fluctuation in temperature data occurred. Ultimately,

the cause is unknown, and the data were left in.

Array 	Julian	Day	Time	Error

150	107	17	1430	Air temp difference from  17 ( 107) 1430 ( 33.519) to

17 ( 107) 1445 ( 36.609) is greater than 3.0 degrees C

Data points were reviewed before and after the following error. There was a

notable change, but the data were consistent.  It was most likely an event.  The

data were left in.

Array 	Julian	Day	Time	Error

150	109	19	1215	Rel hum difference from  19 ( 109) 1215 ( 56.917) to  19

( 109) 1230 ( 95.75) is greater than 25%

Array 	Day	Julian	Time	Error

102	10	100	1900	Wind speed is less than 0.5 m/s from  10 ( 100)  1900

to  11 ( 101)  700

102	15	105	1800	Wind speed is less than 0.5 m/s from  15 ( 105)  1800

to  16 ( 106)  800

102	16	106	1900	Wind speed is less than 0.5 m/s from  16 ( 106)  1900

to  17 ( 107)  800

102	18	108	1800	Wind speed is less than 0.5 m/s from  18 ( 108)  1800

to  19 ( 109)  700

102	21	111	1900	Wind speed is less than 0.5 m/s from  21 ( 111)  1900

to  22 ( 112)  2000

102	25	115	1900	Wind speed is less than 0.5 m/s from  25 ( 115)  1900

to  26 ( 116)  700

102	27	117	1700	Wind speed is less than 0.5 m/s from  27 ( 117)  1700

to  29 ( 119)  800

May 2002

Data points were reviewed before and after the following errors. There was a

notable change, but the data were consistent.  It was most likely an event.  The

data were left in.

Array 	Julian	Day	Time	Error

150	151	31	1315	Air temp difference from  31 ( 151) 1315 ( 24.429) to

31 ( 151) 1330 ( 20.242) is greater than 3.0 degrees C

150	151	31	1715	Air temp difference from  31 ( 151) 1715 ( 20.772) to

31 ( 151) 1730 ( 17.752) is greater than 3.0 degrees C

Data points were reviewed before and after the following errors. The cause is

unknown, and it was most likely an event.  The data were left in.

Array 	Julian	Day	Time	Error

150	141	21	1315	Rel hum difference from  21 ( 141) 1315 ( 58.12) to  21

( 141) 1330 ( 85.336) is greater than 25%

150	141	21	1330	Rel hum difference from  21 ( 141) 1330 ( 85.336) to  21

( 141) 1345 ( 60.198) is greater than 25%

150	141	21	1415	Rel hum difference from  21 ( 141) 1415 ( 51.014) to  21

( 141) 1430 ( 76.353) is greater than 25%

150	141	21	1445	Rel hum difference from  21 ( 141) 1445 ( 91.838) to  21

( 141) 1500 ( 64.019) is greater than 25%

Data points were reviewed before and after the following errors. There was a

notable change, and it was most likely an event.  The data were left in.

Array 	Julian	Day	Time	Error

151	122	2	1930	Precip difference from  31 ( 122) 1930 ( .508) to  31 (

122) 1945 ( 5.588) is greater than 5 mm

151	122	2	1945	Precip difference from  31 ( 122) 1945 ( 5.588) to  31 (

122) 2000 ( .508) is greater than 5 mm

Array 	Day	Julian	Time	Error

102	1	121	1800	Wind speed is less than 0.5 m/s from  1 ( 121)  1800

to  2 ( 122)  900

102	4	124	1800	Wind speed is less than 0.5 m/s from  4 ( 124)  1800

to  5 ( 125)  800

102	5	125	1800	Wind speed is less than 0.5 m/s from  5 ( 125)  1800

to  6 ( 126)  700

102	8	128	1800	Wind speed is less than 0.5 m/s from  8 ( 128)  1800

to  9 ( 129)  900

102	11	131	1800	Wind speed is less than 0.5 m/s from  11 ( 131)  1800

to  13 ( 133)  2300

102	19	139	2000	Wind speed is less than 0.5 m/s from  19 ( 139)  2000

to  20 ( 140)  800

102	22	142	1900	Wind speed is less than 0.5 m/s from  22 ( 142)  1900

to  23 ( 143)  800

102	26	146	1800	Wind speed is less than 0.5 m/s from  26 ( 146)  1800

to  27 ( 147)  1100

102	27	147	1800	Wind speed is less than 0.5 m/s from  27 ( 147)  1800

to  28 ( 148)  1400

102	28	148	1600	Wind speed is less than 0.5 m/s from  28 ( 148)  1600

to  29 ( 149)  1100

The following hourly and daily data were deleted on May 3 due to loss of 5 second data

resulting from station maintenance from 1400-1430.

Array 	Day	Julian	Time	Error

101	3	123	1500	Technician changed 101 Array data from 3 ( 123) 1500

102	3	123	1500	Technician changed 102 Array from 3 ( 123) 1500

241	3	123	2400	Technician changed 241 Array from  3 ( 123) 2400

242	3	123	2400	Technician changed 242 Array from  3 ( 123) 2400

243	3	123	2400	Technician changed 243 Array data from  3 ( 123) 2400

244	3	123	2400	Technician changed 244 Array data from  3 ( 123) 2400

Pressure sensor errors (-99999) were recorded on May 20 @ 1900 in array 150 and @ 2400

in array 244.  Reason unknown.

June 2002

The station was not collecting data due to equipment maintenance on 06/03/02

from 800-1400. The following data were removed:

Array 	Day	Julian	Time	Error

101      3       154      800  Technician changed 101 Array data from 3

( 154) 800 to  3 ( 154) 1300

102      3           154      800  Technician changed 102 Array from 3 (

154) 800 to  3 ( 154) 1300

241      3           154     2400  Technician changed 241 Array from  3 (

154) 2400 to  3 ( 154) 2400

242      3           154     2400  Technician changed 242 Array from  3 (

154) 2400 to  3 ( 154) 2400

243      3           154     2400  Technician changed 243 Array data from  3

( 154) 2400 to  3 ( 154) 2400

244      3           154     2400  Technician changed 244 Array data from  3

( 154) 2400 to  3 ( 154) 2400

Anomalous criteria for wind speed (i.e. "wind speed is less than 0.5m/s") were

listed many times in the WDMP error report for this month as possible errors.

We are certain the sensor was operating properly, most likely there was simply

no wind.  The data were left in.

Data points were reviewed before and after the following errors. Other

parameters changed along with the air temperature, it was most likely an event.

The data were left in.

Array 	Day	Julian	Time	Error

150     22           173     1315  Air temp difference from  22 ( 173) 1315

( 28.418) to  22 ( 173) 1330 ( 25.071) is greater than 3.0 degrees C

150     23           174     1900  Air temp difference from  23 ( 174) 1900

( 28.496) to  23 ( 174) 1915 ( 25.283) is greater than 3.0 degrees C

150     26           177     1415  Air temp difference from  26 ( 177) 1415

( 31.8) to  26 ( 177) 1430 ( 22.773) is greater than 3.0 degrees C

The data were left in.

Array 	Day	Julian	Time	Error

150     26           177     1415  Rel hum difference from  26 ( 177) 1415 (

59.822) to  26 ( 177) 1430 ( 91.751) is greater than 25%

An error message of "Pressure is greater than 1040 or less than 980" occurred

for all 15 minute data from 06/03/02 1400 - 06/30/02 2400. This was due to

improper installation of the barometric pressure sensor (the jumper was not

installed). These data were removed.

Array 	Day	Julian	Time	Error

150      3       154     1400  Technician changed 150 Array data at  3 (154) 1400

to  30 ( 181) 2400

101	3	154	1400	Technician changed 101 Array data at  3 ( 154) 1400

to  30 ( 181) 2400

241	3	154	2400	Technician changed 241 Array data at  3 ( 154) 2400

to  30 ( 181) 2400

243	3	154	2400	Technician changed 243 Array data at  3 ( 154) 2400

to  30 ( 181) 2400

244	3	154	2400	Technician changed 244 Array data at  3 ( 154) 2400

to  30 ( 181) 2400

Array 	Day	Julian	Time	Error

151	5	156	2145	Precip difference from  5 ( 156) 2145 ( 6.096)

to  5 ( 156) 2200 ( 1.016) is greater than 5 mm

151	26	177	1430	Precip difference from  26 ( 177) 1430 ( 5.588)

to  26 ( 177) 1445 ( .254) is greater than 5 mm

July 2002

Anomalous criteria for wind speed (i.e. "wind speed is less than 0.5m/s") were

listed many times in the WDMP error report for this month as possible errors.

We are certain the sensor was operating properly, most likely there was simply

no wind.  The data were left in.

An error message of "Pressure is greater than 1040 or less than 980" occurred

for all 15 minute data from 07/01/02 0000 - 07/31/02 2400. This was due to

improper installation of the barometric pressure sensor (the jumper was not

installed). These data were removed.

Array 	Day	Julian	Time	Error

101	1	182	100	Technician changed 101 Array data at  1 ( 182) 100

to  31 ( 212) 2400

150	1	182	15	Technician changed 150 Array data at  1 ( 182) 15

to  31 ( 212) 2400

241	1	182	2400	Technician changed 241 Array data at  1 ( 182) 2400

to  31 ( 212) 2400

243	1	182	2400	Technician changed 243 Array data at  1 ( 182) 2400

to  31 ( 212) 2400

244	1	182	2400	Technician changed 244 Array data at  1 ( 182) 2400

to  31 ( 212) 2400

Data points were reviewed before and after the following error. Other parameters

changed along with the air temperature, it was most likely an event.  The data

were left in.

Array 	Day	Julian	Time	Error

150     23           204     1300  Air temp difference from  23 ( 204) 1300

( 32.746) to  23 ( 204) 1315 ( 23.848)

Data points were reviewed before and after the following error. Other parameters

changed along with the relative humidity, it was most likely an event.  The data

were left in.

Array 	Day	Julian	Time	Error

150     23           204     1300  Rel hum difference from  23 ( 204) 1300 (

57.488) to  23 ( 204) 1315 ( 89.882) is greater than 25%

August 2002

Anomalous criteria for wind speed (i.e. "wind speed is less than 0.5m/s") were

listed many times in the WDMP error report for this month as possible errors.

We are certain the sensor was operating properly, most likely there was simply

no wind.  The data were left in.

An error message of "Pressure is greater than 1040 or less than 980" occurred

for all 15 minute data from 08/01/02 0000 - 08/30/02 2400. This was due to

improper installation of the barometric pressure sensor (the jumper was not

installed). These data were removed.

Array 	Day	Julian	Time	Error

101	1	213	100	Technician changed 101 Array data at  1 ( 213) 100

to  31 ( 243) 2400

150	1	213	15	Technician changed 150 Array data at  1 ( 213) 15

to  31 ( 243) 2400

241	1	213	2400	Technician changed 241 Array data at  1 ( 213) 2400

to  31 ( 243) 2400

243	1	213	2400	Technician changed 243 Array data at  1 ( 213) 2400

to  31 ( 243) 2400

244	1	213	2400	Technician changed 244 Array data at  1 ( 213) 2400

to  31 ( 243) 2400

Data points were reviewed before and after

the following errors. Many parameters exibited a sudden change while errors were

recorded, it was most likely an event.  All the following the data were left in.

Array 	Day	Julian	Time	Error

150      2           214     1715  Air temp difference from  2 ( 214) 1715 (

23.44) to  2 ( 214) 1730 ( 20.224) is greater than 3.0 degrees C

150     16           228     1400  Air temp difference from  16 ( 228) 1400

( 33.881) to  16 ( 228) 1415 ( 25.248) is greater than 3.0 degrees C

150     18           230     1530  Air temp difference from  18 ( 230) 1530

( 33.41) to  18 ( 230) 1545 ( 27.523) is greater than 3.0 degrees C

Array ID Calendar Date  Julian Day Time   Error Message

150     16           228     1400  Rel hum difference from  16 ( 228) 1400 (

44.636) to  16 ( 228) 1415 ( 84.588) is greater than 25%

150     18           230     1530  Rel hum difference from  18 ( 230) 1530 (

44.635) to  18 ( 230) 1545 ( 76.156) is greater than 25%

Array 	Day	Julian	Time	Error

151      2           214     1715  Precip difference from  2 ( 214) 1715 (

.254) to  2 ( 214) 1730 ( 7.366) is greater than 5 mm

151      2           214     1730  Precip difference from  2 ( 214) 1730 (

7.366) to  2 ( 214) 1745 ( .254) is greater than 5 mm

September 2002

Anomalous criteria for wind speed (i.e. "wind speed is less than 0.5m/s") were

listed many times in the WDMP error report for this month as possible errors.

We are certain the sensor was operating properly, most likely there was simply

no wind.  The data were left in.

An error message of "Pressure is greater than 1040 or less than 980" occurred

for all 15 minute data from 09/01/02 0000 - 09/30/02 2400. This was due to

improper installation of the barometric pressure sensor (the jumper was not

installed). These data were removed.

Array 	Day	Julian	Time	Error

101	1	244	100	Technician changed 101 Array data at  1 ( 244) 100

to  30 ( 273) 2400

150	1	244	15	Technician changed 150 Array data at  1 ( 244) 15

to  30 ( 273) 2400

241	1	244	2400	Technician changed 241 Array data at  1 ( 244) 2400

to  30 ( 273) 2400

243	1	244	2400	Technician changed 243 Array data at  1 ( 244) 2400 to  30 ( 273) 2400"

244	1	244	2400	Technician changed 244 Array data at  1 ( 244) 2400

to  30 ( 273) 2400

Data points were reviewed before and after the following errors. The relative

humidity was near 100%, it was most likely a rain event.  The data were left in.

Array 	Day	Julian	Time	Error

151     27           270     2130  Precip difference from  27 ( 270) 2130 (

1.27) to  27 ( 270) 2145 ( 9.906) is gr

151     27           270     2145  Precip difference from  27 ( 270) 2145 (

9.906) to  27 ( 270) 2200 ( 4.318) is g

October 2002

Anomalous criteria for wind speed (i.e. "wind speed is less than 0.5m/s") were

listed many times in the WDMP error report for this month as possible errors.

We are certain the sensor was operating properly, most likely there was simply

no wind.  The data were left in.

An error message of "Pressure is greater than 1040 or less than 980" occurred

for all 15 minute data from 10/01/02 0000 - 10/31/02 2400. This was due to

improper installation of the barometric pressure sensor (the jumper was not

installed). These data were removed.

Array 	Day	Julian	Time	Error

101	1	274	100	Technician changed 101 Array data at  1 ( 274) 100

to  31 ( 304) 2400

150	1	274	15	Technician changed 150 Array data at  1 ( 274) 15

to  31 ( 304) 2400

241	1	274	2400	Technician changed 241 Array data at  1 ( 274) 2400

to  31 ( 304) 2400

243	1	274	2400	Technician changed 243 Array data at  1 ( 274) 2400

to  31 ( 304) 2400

244	1	274	2400	Technician changed 244 Array data at  1 ( 274) 2400

to  31 ( 304) 2400

Data points were reviewed before and after the following error. There was a

change in the readings, but the cause is unknown. The data were suspect, but

left in.

Array 	Day	Julian	Time	Error

150     27           300      500  Air temp difference from  27 ( 300) 500 (

4.5636) to  27 ( 300) 515 ( 7.5838) is greater than 3.0 degrees C

November 2002

Anomalous criteria for wind speed (i.e. "wind speed is less than 0.5m/s") were

listed many times in the WDMP error report for this month as possible errors.

We are certain the sensor was operating properly, most likely there was simply

no wind.  The data were left in.

An error message of "Pressure is greater than 1040 or less than 980" occurred

for all 15 minute data from 11/01/02 0000 - 11/07/02 1030. This was due to

improper installation of the barometric pressure sensor (the jumper was not

installed). These data were removed.

Array 	Day	Julian	Time	Error

101	1	305	100	Technician changed 101 Array data from 1 ( 305) 100

to  7 ( 311) 1100

150	1	305	15	Technician changed 150 Array data from  1 ( 305) 15

to  7 ( 311) 1030

241	1	305	2400	Technician changed 241 Array from  1 ( 305) 2400 to

6 ( 310) 2400

243	1	305	2400	Technician changed 243 Array data from  1 ( 305) 2400

to  6 ( 310) 2400

244	1	305	2400	Technician changed 244 Array data from  1 ( 305) 2400

to  7 ( 311) 2400

Data points were reviewed before and after the following error. The data appears

consistent, and it was likely an event. The data were left in.

Array 	Day	Julian	Time	Error

150     11           315     1330  Air temp difference from  11 ( 315) 1330

( 21.162) to  11 ( 315) 1345 ( 18.081) is greater than 3.0 degrees C

The following error occurred because the "1 hour data" is averaged from 5-second

readings taken during that hour while the "15 minute data" (in this case the 15

minute maximum) is not averaged, and only recorded at 15-minute intervals.  The

data were left in.

Array 	Day	Julian	Time	Error

101     16           320     1300  Air temp average in 1 hour data (-.4325)

is less than 15 minute minimum  (-.38942) by at least 10%

101     16           320     1500  Air temp average in 1 hour data (-.22167)

is less than 15 minute minimum (-.18801) by at least 10%

101     27           331      400  Air temp average in 1 hour data ( .02151)

is greater than 15 minute maximum ( .00669) by at least 10%

December 2002

Data points were reviewed before and after the following error. The data appears

consistent, and it was likely an event. The data were left in.

Array 	Day	Julian	Time	Error

150     19           353     1100  Air temp difference from  19 ( 353) 1100

(-1.3718) to  19 ( 353) 1115 ( 3.3897) is greater than 3.0 degrees C

Anomalous criteria for wind speed (i.e. "wind speed is less than 0.5m/s") were

listed a few times in the WDMP error report for this month as possible errors.

We are certain the sensor was operating properly, most likely there was simply

no wind.  The data were left in.

Array 	Day	Julian	Time	Error

102      4           338     2000  Wind speed is less than 0.5 m/s from  4 (

338)  2000 to  5 ( 339)  1200

102      5           339     1400  Wind speed is less than 0.5 m/s from  5 (

339)  1400 to  6 ( 340)  1800

102     10           344     1500  Wind speed is less than 0.5 m/s from  10

( 344)  1500 to  11 ( 345)  1000

102     12           346     1300  Wind speed is less than 0.5 m/s from  12

( 346)  1300 to  14 ( 348)  600

102     15           349     1800  Wind speed is less than 0.5 m/s from  15

( 349)  1800 to  16 ( 350)  900

102     19           353     2400  Wind speed is less than 0.5 m/s from  19

( 353)  2400 to  20 ( 354)  1300

102     24           358     1600  Wind speed is less than 0.5 m/s from  24

( 358)  1600 to  25 ( 359)  700

102     27           361     1400  Wind speed is less than 0.5 m/s from  27

( 361)  1400 to  28 ( 362)  1000

102     28           362     1600  Wind speed is less than 0.5 m/s from  28

( 362)  1600 to  29 ( 363)  1200

The following error occurred because the "1 hour data" is averaged from 5-second

readings taken during that hour while the "15 minute data" (in this case the 15

minute maximum) is not averaged, and only recorded at 15-minute intervals.  The

data were left in.

Array 	Day	Julian	Time	Error

101      7           341     1800  Air temp average in 1 hour data (-.02553)

is greater than 15 minute maximum (-.05381) by at least 10%

101     20           354     2400  Air temp average in 1 hour data ( .53791)

is greater than 15 minute maximum ( .47657) by at least 10%

Deleted erroneous daily minimum barometric pressure value on December 20 from array

244.

Array 	Day	Julian	Time	Error

244	20	354	2400	Technician changed 244 Array data from  20 ( 354) 2400

12. Missing data:

Arrays:

During 2022 all pre-2007 weather data were revisited by the CDMO. Historically those datasets included 15 minute, hourly (60), and daily data arrays (144). As directed by the NERRS Data Management Committee, the CDMO removed the hourly and daily data arrays leaving only the 15 minute data to make the entire NERRS SWMP weather dataset consistent in its reporting. All references to the 60 and 144 arrays were left in the metadata document as they may still provide valuable information, but users should be aware that they are largely no longer relevant. The updated datasets were uploaded to the database and made available through the various data applications at www.nerrsdata.org/get/landing.cfm throughout the fall of 2022.

Data are missing due to equipment failure, or no probes deployed, or maintenance

of equipment.  To find out more details about missing data, contact the Research

Coordinator at the site submitting the data.

January 2002

The equipment was under repair from 01/03/02 1300 - 01/14/02 1630. Data is

missing for this time period.

Array 	Julian	Day	Time	Error

150	3	3	1315	Missing 150 Array data (15 minute data) from

3 ( 3) 1315 to  14 ( 14) 1615

101	3	3	1400	Missing 101 Array data (Hourly Averages) from

3 ( 3) 1400 to  14 ( 14) 1600

102	3	3	1400	Missing 102 Array data (Hourly Average Wind Parameters)

from  3 ( 3) 1400 to  14 ( 14) 1600

241	3	3	2400	Missing 241 data (Daily Averages) from  3 ( 3) 2400

to  13 ( 13) 2400

242	3	3	2400	Missing 242 data (Daily Average Wind Parameters)

from  3 ( 3) 2400 to  13 ( 13) 2400

243	3	3	2400	Missing 243 data (Daily Max/Time Values) from

3 ( 3) 2400 to  13 ( 13) 2400

244	3	3	2400	Missing 244 data (Daily Min/Time Values) from

3 ( 3) 2400 to  13 ( 13) 2400

May 2002

Data was missing May 3 1415-1430 due to maintenance on the station.

Array 	Julian	Day	Time	Error

150	123	3	1415	Missing 150 Array data (15 minute data) from  3 ( 123)

1415 to  3 ( 123) 1430

June 2002

Array 	Day	Julian	Time	Error

150      3      154      800  Missing 150 Array data (15 minute data)

from  3 ( 154) 800 to  3 ( 154) 1345

101      3      154      700  Missing 101 Array data (Hourly Averages)

from  3 ( 154) 900 to  3 ( 154) 2300

102      3      154      700  Missing 102 Array data (Hourly Average

Wind Parameters) from  3 ( 154) 900 to  3 ( 154) 2300

13. Other Remarks/notes:

On 10/10/2023 this dataset was updated to include embedded QAQC flags for anomalous/suspect data.  System-wide monitoring data beginning in 2007 were processed to allow for QAQC flags and codes to be embedded in the data files rather than detailed in the metadata alone (as in the anomalous/suspect, deleted, and missing data sections above).  Prior to 2007, rejected data were deleted from the dataset so they are unavailable to be used at all, but suspect data were only noted in the metadata document.  Suspect data flags <1> were embedded retroactively in order to allow suspect data to be easily identified and filtered from the dataset if desired for analysis and reporting purposes.  No other flags or codes were embedded in the dataset and users should still refer to the detailed explanations above for more information.

Arrays:

During 2022 all pre-2007 weather data were revisited by the CDMO. Historically those datasets included 15 minute, hourly (60), and daily data arrays (144). As directed by the NERRS Data Management Committee, the CDMO removed the hourly and daily data arrays leaving only the 15 minute data to make the entire NERRS SWMP weather dataset consistent in its reporting. All references to the 60 and 144 arrays were left in the metadata document as they may still provide valuable information, but users should be aware that they are largely no longer relevant. The updated datasets were uploaded to the database and made available through the various data applications at www.nerrsdata.org/get/landing.cfm throughout the fall of 2022.

Precipitation:

During the initial years of NERRS SWMP weather data collection the CR10X programming was inconsistent in how precipitation values were recorded. For most reserves, zeros were not recorded when rainfall had not occurred between 2001-2003, instead no rainfall was represented by a blank cell. The CDMO verified which datasets were impacted by this issue for the 2001-2006 datasets and inserted zeros when the metadata indicated that no precipitation occurred and data were not missing for other reasons. In some cases, zero values for precipitation data were evaluated and removed where the metadata confirmed that no rainfall should have been in the dataset. The pre-2007 data did not go through a thorough QAQC process again at that time (in addition to previous QAQC); however, if discrepancies were noticed between what was documented in the metadata and what was in the dataset, additional updates may have been made. The updated datasets were uploaded to the database and made available through the various data applications at www.nerrsdata.org/get/landing.cfm throughout the fall of 2022.

PAR flagging update August 2024:

During SWMP data synthesis work it was noted that PAR values were anomalously high beginning in 2001 and returned to more expected values following the sensor swap on June 3, 2002. PAR data flagging was changed by the CDMO to suspect from January 10, 2001 15:15 through June 3, 2002 07:45.

In June 2009, in order to repopulate data tables, the Centralized Data Management Office removed all -99999 from SWMP weather data files and replaced them with -99.

The Centralized Data Management Office converted all SWMP weather data collected with CR10X

program versions prior to version 4.0 which was distributed in October 2003.  This was necessary

in order to merge the old data format (12 array output) with the new data format found in version 4.0

(3 array output).  The new format produces averages, maximums and minimums every fifteen minutes

(array 15), every hour (array 60) and every day (array 144) for any sensors hooked up to the CR10X.

Specifically, the 150 and 151 fifteen minute data were converted to the new 15 array;

the hourly 101, 102, 105 and 106 data were converted to the new 60 array; and the daily 241, 242,

243, 244, 245 and 246 data were converted to the new 144 array.  With the new format, the use of

55555's to code for deleted data and 11111's to code for missing data has been abandoned.

Hence, all 55555's or 11111's contained in the SWMP weather data collected prior to Version 4.0

of the CR10X program were removed and left blank.

All sensors were replaced with newly calibrated sensors on 06/03/02 1400.

The following data are Daily Totals for Precipitation:  (Note that there are no monthly totals during months where there was missing data).

Jan          2002

Calendar Day         Daily Precipitation Totals (mm)

17			4.064

20			1.270

21			.254

22			4.572

23			.762

24			.508

30			2.794

31			3.810

Feb       2002

Calendar Day      Daily Precipitation Totals (mm)

1             		7.620

10             		7.874

11            		 2.032

17             		3.048

20             		1.524

22              		.254

27             		1.524

Monthly Tot   23.9

Mar     2002

Calendar Day      Daily Precipitation Totals (mm)

2			2.540

3			15.748

9			.254

10			3.556

16			8.636

18			4.318

20			18.542

26			20.574

27			1.270

30			4.318

31			2.794

"Monthly Total"	82.5

Apr          2002

Calendar Day         Daily Precipitation Totals (mm)

1			2.540

3			7.874

9			.762

12			.508

13			4.318

14			3.810

15			5.334

19			3.302

20			.762

22			2.540

25			13.716

28			33.528

29			2.286

30			1.016

"Monthly Total"	82.3

May         2002

Calendar Day         Daily Precipitation Totals (mm)

2			19.558

3			.254

7			.508

9			.508

12			9.144

13			33.274

14			4.318

17			.254

18			23.622

21			1.016

24			.254

26			3.048

28			1.270

31			13.208

Jun          2002

Calendar Day         Daily Precipitation Totals (mm)

5               		12.192

6               		23.622

7               		10.922

12                 		.508

14               		13.462

15               		12.446

16                		1.778

18                		1.016

22                		8.636

23                		1.524

24                 		.254

26                		6.604

27                		1.778

Jul          2002

Calendar Day         Daily Precipitation Totals (mm)

9			2.794

18			.254

19			6.604

20			.254

23			9.144

24			.254

27			.254

28			.508

"Monthly Total"	20.1

Aug          2002

Calendar Day         Daily Precipitation Totals (mm)

1                		3.810

2                		9.398

3                 		.254

5                		9.906

16                		1.270

18                 		.508

20               		12.446

24                		4.572

25                 		.254

28                 		.254

29               		34.798

Monthly Total     77.5

Sep          2002

Calendar Day         Daily Precipitation Totals (mm)

1                 		.254

4                		6.096

15               		21.336

16               		18.288

17                 		.254

22                		4.826

26                		8.382

27               		38.354

28                 		.508

Monthly Total     98.3

Oct          2002

Calendar Day         Daily Precipitation Totals (mm)

2                		1.016

4                		4.064

5                		2.032

7                 		.508

10                		2.794

11               		39.624

12               		38.862

13                		1.524

16               		27.432

17                		1.270

19                 		.508

25                		3.556

26               		17.018

Monthly Total    140.2

Nov          2002

Calendar Day         Daily Precipitation Totals (mm)

4                 		.254

5                		3.302

6               		14.986

11                		2.032

12               		17.272

13                 		.762

16               		19.558

17               		24.638

18                		7.620

19                 		.508

21                		2.032

22               		14.986

23                 		.254

26                 		.254

27                		2.794

Monthly Total    111.3

Dec          2002

Calendar Day         Daily Precipitation Totals (mm)

5                		5.588

11                		8.128

12               		30.480

13                		6.350

14               		10.668

16                 		.254

20               		16.002

25               		14.224

26               		12.954

30                 		.254

31                		2.032

Monthly Total    106.9