Hudson River (HUD) NERR Water Quality Metadata

January 1, 2023 – December 2023

Latest Update: 07/15/2025

I.  Data Set and Research Descriptors

1)  Principal investigator(s) and contact persons – 								Chris Mitchell, Research Coordinator

Email: Christopher.Mitchell@dec.ny.gov

Phone: (845) 889-4745 x119

Fax: (845) 889-4749

Christina Pacella, SWMP Technician

Email: Christina.Pacella@dec.ny.gov

Phone: (845) 889-4745 x104

Address:

Hudson River NERR

Norrie Point Environmental Center

256 Norrie Point Way

PO Box 315

Staatsburg, NY  12580

Fax: (845) 889-4749

2)  Entry verification –

Deployment data are uploaded from the YSI data logger to a personal computer with Windows 7 or newer operating system.  Files are exported from KOR Software in a comma separated file (CSV) and uploaded to the CDMO where they undergo automated primary QAQC; automated Depth/Level corrections for changes in barometric pressure (cDepth or cLevel parameters); and become part of the CDMO’s online provisional database.  All pre- and post-deployment data are removed from the file prior to upload.  During primary QAQC, data are flagged if they are missing or out of sensor range.  The edited file is then returned to the reserve for secondary QAQC where it is opened in Microsoft Excel and processed using the CDMO’s NERRQAQC Excel macro.  The macro inserts station codes, creates metadata worksheets for flagged data and summary statistics, and graphs the data for review.  It allows the user to apply QAQC flags and codes to the data, remove any overlapping deployment data, append files, and export the resulting data file for upload to the CDMO.  Upload after secondary QAQC results in ingestion into the database as provisional plus data, recalculation of cDepth or cLevel parameters, and finally tertiary QAQC by the CDMO and assimilation into the CDMO’s authoritative online database.  Where deployment overlap occurs between files, the data produced by the newly calibrated sonde is generally accepted as being the most accurate.  For more information on QAQC flags and codes, see Sections 11 and 12.

Persons responsible for data verification are the Research Assistant.

3)  Research objectives –

The Hudson River National Estuarine Research Reserve (HRNERR; the Reserve) protects nearly 5,000 acres of diverse tidal and upland habitats across four distinct sites within the Hudson River Estuary (the Estuary). The Estuary is approximately 153 miles long and is profoundly influenced by the ocean’s tides. Four component sites comprise the Reserve (from north to south): Stockport Flats, Tivoli Bays, Iona Island, and Piermont Marsh. Each of these component sites contain different kids and combinations of tidal habitats, and together they represent gradients found within the Estuary, including salinity and degree of watershed development.

Surface water is monitored at each of the component sites utilizing YSI EXO2 multiparameter sondes; descriptions of each station are described in Section 4. These sondes monitor the ebbing and flooding Hudson River water.  Examining the influence of tidal exchange allows identification of long-term trends in the water quality of the Hudson River Estuary at these locations.

4)  Research methods –

Historic Sites

YSI 6600’s (6600 EDS starting 2003) were deployed at the Hudson River NERR Component sites from 1996 – 2013.  For information regarding the previous deployments, please refer to the metadata from prior years.

Sawkill (SK; 1995-2019)

From September 2011 to December 2013 the Sawkill deployment tube was affixed to the undamaged end of the dam at a 45-degree angle, retaining the original 0.5 meters off the bottom of the creek.  The station was destroyed again during the winter of 2013-2014.  The station was rebuilt as vertically mounted in a four – inch PVC tube in April 2014.  The sonde was located 0.5 meters from the creek bed and was reconstructed as close as possible to original sampling location.

Stony Creek (SC; 2002-2022)

At Stony Creek, instruments are deployed in a perforated four-inch PVC pipe set in approximately 70 pounds of concrete at such angle that when laid in a creek pool, the streambank-end of the pipe matched the grade of the bank for most of its length and the sonde, set in place by a 0.5 inch through-bolt, rests 0.25 meters above the substrate.  Along the portion of the PVC in contact with dry bank, several rebar stakes, bent around and lashed through holes in the pipe, hold the deployment in place and discourage vandalism.  Additionally, the tube is painted to match surrounding deadfall, the end is capped, and locked and contact/project information provided to further discourage tampering.

Current Sites

The following describes the five locations in which EXO2 water quality monitoring instruments are deployed. These sites are listed from north to south along the Hudson River Estuary.

Ferry Landing (FL; 2022 – Current)

The instrument at Ferry Landing was established in March of 2022 and is in a 4” vertical PVC pipe attached to a steel piling what was driven to support a fishing pier. The instrument rests on a horizontal stainless-steel bolt which is 0.25 m above the substrate.  This instrument must be pulled during the winter months, and the shallow nature of the locale and its propensity to freeze through.

Tivoli North Bay (TN; 1999-Current)

A vertical four-inch PVC pipe is lashed to wooden pilings under the railroad culvert in Tivoli North Bay.  The sonde is secured within this pipe and situated 0.5m from the substrate, resting on a 0.5-inch stainless steel bolt.

Tivoli South Bay (TS; 1995-Current)

The instrument in Tivoli South Bay is deployed 0.5 meters off the bottom in a perforated four-inch PVC tube that is vertically mounted to existing concrete and metal structures. The instrument at Tivoli South Bay historically used a depth sensor that was vented to the atmosphere. Beginning from the 2014 sampling season, the instrument now utilizes a non-vented sensor for depth readings.

Norrie Point (NP; 2008-Current)

A “Secondary SWMP” station was established at the reserve’s Norrie Point, office in 2018. This station was originally constructed in June 2008, an operated independently of SWMP until 2018.  Historical data are available upon request.

The instrument at this location is identical to those mentioned above for primary SWMP use. The instrument is deployed in a ported 4” aluminum pipe, which is lashed to a large wooden piling structure.  The sonde is deployed vertically in the pipe and sits on a stainless-steel bolt 0.5 meters above the river bottom. Due to the extra depth and protection of the Norrie Point station, the sonde is deployed year-round.  Dataloggers are calibrated and swapped from March through December of each year. The instrument does become “locked” in place during the ice months (December to March), and typically cannot be accessed during this time.  It is retrieved as soon as ice-out occurs.

Bear Mountain (BM; 2020 – Current)

The sonde at Bear Mountain is deployed in a vertical 4-inch PVC pipe secured to a wooden dock structure on the south side of the dock. The sonde rests 0.5 m off the river bottom on a 0.5-inch bolt. This site was initially built in 2019 and brought on for primary SWMP use in 2020. It is deployed year-round as its typical depth and location near the salt front does not typically the instrument to be impacted by ice.

Methodologies

Initially, data were collected at 30-minute intervals.  The interval was changed to 15 minutes on the 08/03/2006 deployment at Tivoli South, and on 08/09/2006 at Stony Creek, Saw Kill Creek, and Tivoli North.  A 15-minute interval was maintained through the current data set.  Each YSI EXO 2 datalogger measures depth, dissolved oxygen, temperature, conductivity, salinity, pH, and turbidity.  In addition, from 10/07/2009, chlorophyll is measured at Tivoli North and South locations.  As of 2014 all sampling locations measure chlorophyll.  Prior to deployment, calibration and maintenance are performed on each datalogger following the manufacturer's instructions (YSI EXO operations manual).  Calibration standards are utilized for pH, turbidity, and conductivity.  These standards are purchased from a scientific supply company. Standard Lot number is recorded before each calibration.  Chlorophyll data were calibrated using a one-point calibration until August of 2022. The one-point calibration utilized deionized (DI) water and a zero-point calibration. The YSI EXO operations manual and NERRS SWMP EXO Manual both call for a two-point calibration. As of August 20223, a two-point calibration has been implemented for chlorophyll sensors.

Individual instrument deployments last from two to six weeks. At the end of each deployment period, the dataloggers are swapped with newly calibrated instruments to avoid breaks in data collection.  After instruments are retrieved, calibration is checked against standards and the probes are serviced.  The data are uploaded to a desktop computer via Bluetooth. Graphs of each parameter are generated and examined immediately to identify significant problems or events that may have occurred during deployment. The data are then verified as described previously.

A Sutron Sat-Link2 transmitter was installed at the Tivoli South station on 11/15/2005 and transmits data to the NOAA GOES satellite, NESDIS ID #3B00A782.  A second Sutron Sat-Link2 transmitter was also installed at the Sawkill station in June 2012 and transmitted data to the NOAA GOES satellite, NESDIS ID #3B010580.  The Sawkill station was discontinued after the 2019 season.  The Bear Mountain station now utilizes its telemetry equipment and occupies NESDIS ID #3B010580.  For the 2014 season, a radio and Satlink 2 combination had been installed at the Stony Creek sampling location and transmitted data to the NOAA GOES satellite, NESDIS ID #3B03A08C.  The Stony Creek station was discontinued following the 2021 sampling season.  The Ferry Landing station now utilizes its telemetry equipment and occupies NESDIS ID #3B03A08C. Beginning with the 2018 sampling season at Norrie Point, a Sat-Link2 transmitter transmits data to the NOAA GOES satellite, NESDIS ID # 3B0467A6. On 10/6/2023 the Sat-Link2 transmitter at the Norrie Point was upgraded to a STORM-3 data logger and utilizes the same NESDIS ID.  The transmissions are scheduled hourly and contain four (4) data sets reflecting 15-minute data sampling intervals.  Upon receipt by the CDMO, the data undergoes the same automated primary QAQC process detailed in Section 2 above.  The “real-time” telemetry data become part of the provisional dataset until undergoing secondary and tertiary QAQC and assimilation in the CDMO’s authoritative online database.  Provisional and authoritative data are available at http://cdmo.baruch.sc.edu.

5)  Site location and character –

The research reserve spans the middle 100 miles of the Hudson River estuary and is comprised of four sites: Stockport Flats, Tivoli Bays, Iona Island, and Piermont Marsh.  The entirety of the 150-mile tidal portion of the river, experiences on average a 1.5-meter tidal range, until it reaches an engineered head of tide at the Federal Dam in Troy, NY.  All mainstem monitoring stations listed below (excluding SC) record similar tidal amplitudes, although the depth of the monitoring devices varies as listed below.  A salinity gradient exists with freshwater prevalent primarily north of the Iona Island (BM) component site.  Environmental factors such as drought and excessive rainfall may allow this salt gradient to migrate both north and south.  Typical salinity is 0.1 PSU at the northern sites (TN, TS, and NP) and ranging from 0.1 – 10.0 PSU at the southern monitoring location (BM).

Freshwater dominates all but the lower 40 river miles of the Hudson River.  The river flows for a total of 315 miles, with the only the lower 153 miles classified as an estuary.  The upper portion of the river drains a total of 4,590 square miles spanning 3 states and majority of the Adirondack Mountains of New York State.  A significant freshwater input from the Mohawk River, with its confluence located just north of the head of tide, provides a significant source of additional freshwater. (Figure – 1A)

Figure 1-A

| Site name | Ferry Landing (FL) |
| --- | --- |
| Latitude and longitude | 42°21'14.36"N, 73°47'20.76"W |
| Tidal range (meters) | 1.19 m |
| Salinity range (psu) | 0 – 0.20 psu |
| Type and amount of freshwater input | An unnamed tributary is located approximately 1.12 kilometers (km) south of the station which drains approximately 8.9 km2. Stockport Creek is approximately 5km south of the site and has a watershed of roughly 1300 km2. |
| Water depth (meters, MLW) | This information is pending – surveying activities are being conducted during the field season of 2024, and MLW will be updated on a site-by-site basis as it becomes available. |
| Sonde distance from bottom (meters) | 0.25 m |
| Bottom habitat or type | Soft silt and clay |
| Pollutants in area | The Hudson River from Hudson Falls (Washington County) to the Battery in New York City is an Environmental Protection Agency (EPA) superfund site with a classification of 02 for polychlorinated biphenyls in the sediment. This site is outside of any determined hot spots where remediation efforts have been attempted. However, as of Spring 2023 the EPA is undertaking additional investigations in the lower Hudson to evaluate contaminants and further steps. Other potential inputs include combined sewer overflows (CSOs) and runoff from agricultural lands. |
| Description of watershed | This station is situated along the eastern bank of the Hudson River and is affixed to a fishing pier. Land use within the immediate vicinity is comprised predominantly of vegetated wetland and forested areas with some residential development |

| Site name | Tivoli North Bay (TN) |
| --- | --- |
| Latitude and longitude | 42° 02' 11.56464" N, 73° 55' 31.16645" W |
| Tidal range (meters) | 1.19 m |
| Salinity range (psu) | 0 – 0.20 psu |
| Type and amount of freshwater input | Stony Creek is the primary tributary which contributes freshwater input into this site; it drains approximately 35.7 km2 into TN. |
| Water depth (meters, MLW) | -0.57917 m (NAVD88) |
| Sonde distance from bottom (meters) | 0.50 m |
| Bottom habitat or type | Soft silt and clay |
| Pollutants in area | The Hudson River from Hudson Falls (Washington County) to the Battery in New York City is an Environmental Protection Agency (EPA) superfund site with a classification of 02 for polychlorinated biphenyls in the sediment. This site is outside of any determined hot spots where remediation efforts have been attempted. However, as of Spring 2023 the EPA is undertaking additional investigations in the lower Hudson to evaluate contaminants and further steps. Other potential inputs include combined sewer overflows (CSOs) and runoff from agricultural lands. |
| Description of watershed | TN is a freshwater tidal marsh with emergent marsh vegetation dominated by narrowleaf cattail (Typha angustifolia). Located on the eastern side of the Hudson River, it is constrained to the west by a railroad. East of the site are wooded clay bluffs and forested uplands and fields, with some land utilized by Bard College facilities and properties |

| Site name | Tivoli South Bay (TS) |
| --- | --- |
| Latitude and longitude | 42° 01' 37.336" N, 73° 55' 33.445" W |
| Tidal range (meters) | 1.19 m |
| Salinity range (psu) | 0 – 0.20 psu |
| Type and amount of freshwater input | The Sawkill is the primary tributary which contributes freshwater input into this site; it drains approximately 35.4 km2 into TS. |
| Water depth (meters, MLW) | -0.57391 m (NAVD88) |
| Sonde distance from bottom (meters) | 0.50 m |
| Bottom habitat or type | Soft silt and clay |
| Pollutants in area | The Hudson River from Hudson Falls (Washington County) to the Battery in New York City is an Environmental Protection Agency (EPA) superfund site with a classification of 02 for polychlorinated biphenyls in the sediment. This site is outside of any determined hot spots where remediation efforts have been attempted. However, as of Spring 2023 the EPA is undertaking additional investigations in the lower Hudson to evaluate contaminants and further steps. Other potential inputs include combined sewer overflows (CSOs) and runoff from agricultural lands |
| Description of watershed | TS is a freshwater tidal wetland with intertidal mudflats exposed at low tide. During the growing season (June – September), the subtidal area of Tivoli South Bay is dominated by the invasive floating macrophyte water chestnut (Trapa natans). Similar to TN, the surrounding areas are characterized by wooded uplands and Bard College facilities and properties. This site is also abutted to the west by a railroad. |

| Site name | Norrie Point (NP) |
| --- | --- |
| Latitude and longitude | 41°49'54.0"N 73°56'31.0"W |
| Tidal range (meters) | 1.19 m |
| Salinity range (psu) | 0 – 0.20 psu |
| Type and amount of freshwater input | The Indian Kill (known locally as the Enderkill) is a small freshwater tributary located approximately 0.5 km to the north of the station. According to USGS StreamStats, it has a drainage area of approximately 4.7 square miles |
| Water depth (meters, MLW) | This information is pending – surveying activities are being conducted during the field season of 2024, and MLW will be updated on a site-by-site basis as it becomes available. |
| Sonde distance from bottom (meters) | 0.50 m |
| Bottom habitat or type | Soft silt and clay |
| Pollutants in area | The Hudson River from Hudson Falls (Washington County) to the Battery in New York City is an Environmental Protection Agency (EPA) superfund site with a classification of 02 for polychlorinated biphenyls in the sediment. This site is outside of any determined hot spots where remediation efforts have been attempted. However, as of Spring 2023 the EPA is undertaking additional investigations in the lower Hudson to evaluate contaminants and further steps. Other potential inputs include combined sewer overflows (CSOs) and runoff from agricultural lands |
| Description of watershed | The station is located on the east side of the Hudson River at the end of a dock extending out into the main channel. There is a bed of invasive Trapa natans in an adjacent cove during the growing season. The station is within the Mills-Norrie State Park, which features wooded uplands, forested wetlands and vernal pools, a campground, access roads, and recreational trails. Beyond are residential and sparse commercial properties along State Route 9. |

| Site name | Bear Mountain (BM) |
| --- | --- |
| Latitude and longitude | 41° 18' 51.0” N, 73° 59' 6.0" W |
| Tidal range (meters) | 1.19 m |
| Salinity range (psu) | 0.10 - 10 |
| Type and amount of freshwater input | An unnamed stream located approximately 100 m west of the sonde station drains roughly 0.07 km2 into the Hudson River; this stream contributes a slight freshwater signature to the site. |
| Water depth (meters, MLW) | -0.2613 m (NAVD88) |
| Sonde distance from bottom (meters) | 0.50 m |
| Bottom habitat or type | Rocky substrate composed of cobbles and gravel |
| Pollutants in area | The Hudson River from Hudson Falls (Washington County) to the Battery in New York City is an Environmental Protection Agency (EPA) superfund site with a classification of 02 for polychlorinated biphenyls in the sediment. This site is outside of any determined hot spots where remediation efforts have been attempted. However, as of Spring 2023 the EPA is undertaking additional investigations in the lower Hudson to evaluate contaminants and further steps. Other potential inputs include combined sewer overflows (CSOs) and runoff from agricultural lands |
| Description of watershed | BM is situated approximately 700 m north of the Iona Island Marsh component site at River Mile (RM) 45. This is a tidal brackish wetland with emergent vegetation dominated by narrowleaf cattail (Typha angustifolia) and invasive common reed (Phragmites australis). A freshwater creek, Doodletown Brook, is the main tributary flowing into Iona Island Marsh. BM sits at the foot of Bear Mountain State Park, a mostly undeveloped, mountainous woodland. The Bear Mountain samples are frequently in the northern extent of the salt front in the Hudson River. Salinity concentrations at this site are driven primarily by tidal and meteorological factors. |

| Station Code | SWMP Status | Station Name | Location | Active Dates | Reason Decommissioned | Notes |
| --- | --- | --- | --- | --- | --- | --- |
| FL | P | Ferry Landing | 42°21’14.43” N 73°47’20.72"W | 04/01/2022 - Current | NA | NA |
| TN | P | Tivoli North Bay | 42° 2' 11.56 N, 73° 55' 31.17 W | 07/01/1999 00:00 - Current | NA | NA |
| TS | P | Tivoli South Bay | 42° 1' 37.34 N, 73° 55' 33.45 W | 05/01/1995 00:00 - Current | NA | NA |
| NP | S | Norrie Point | 41°49'54.0"N 73°56'31.0"W | (06/27/2008) *01/01/2018 00:00 - Current | NA | *Secondary SWMP status confirmed as of 01/01/2018, prior data may be available per request |
| BM | P | Bear Mountain | 41°18’53.19” N 73°59'6.78"W | 01/01/2020 00:00 - Current | NA | New site established for the 2020 sampling season. Replaced Saw Kill (SK) as primary station in 2019 |
| SK | D | Sawkill | 42° 1' 1.82 N, 73° 54' 53.86 W | 05/01/1995 00:00 –12/3/2019 | The attachment point (dam) is scheduled for demolition | NA |
| SC | D | Stony Creek | 42° 2' 46.68 N, 73° 54' 38.88 W | 04/01/2002 00:00 –12/31/2022 | Site Profile has been significant altered due to slope failure, beaver activity, and anthropogenic changes and no longer is a viable location to deploy instruments. | NA |

6)  Data collection period –

YSI 6-series datalogger sampling at Sawkill (SK) and Tivoli South Bay (TS) began in May 1995. Sampling at Tivoli North Bay (TN) began in July 1996 but ceased in 1997-1998 due to instrument exposure at low tide. In 1999, an alternate site was established for TN and sampling began again in July of 1999.  Sampling began at Stony Creek (SC) in April 2002 and ceased in 2021 due to changes in the site character (see above chart).  Sampling is continuous from April through December but is not conducted from approximately mid-December through the end of March due to winter storms and ice on the Hudson River.

A YSI 6600 Extended Deployment System (EDS) sonde was deployed at Tivoli South Bay starting on 04/08/2003.  The EDS differed from the other 6600 instruments in that a large wiper was positioned centrally to the probes and wiped them prior to every sample. This wiper required that the pH probe did not have a guard over the glass bulb, and a new flat pH probe was deployed on the EDS starting 09/25/2003. Otherwise, the EDS sonde was identical to the other 6600 dataloggers. YSI 6600 V2s had been utilized at Tivoli South Bay and Tivoli North Bay since the addition of chlorophyll probes on 10/07/2009.

All component sites were switched to YSI EXO 2 sondes during the 2014 sampling season. All sites now contain non-vented equipment.

Dates and times for site deployments and retrievals for 2023:

Site: FL

| Deployment | Retrieval |
| --- | --- |
| NO INSTRUMENT DEPLOYED DUE TO ICE |  |
| 03/09/2023 @ 10:45 | 04/20/2023 @ 09:15 |
| 04/20/2023 @ 10:00 | 05/19/2023 @ 09:45 |
| 05/19/2023 @ 10:00 | 06/20/2023 @ 10:45 |
| 06/20/2023 @ 11:00 | 07/24/2023 @ 10:45 |
| 07/24/2023 @ 11:15 | 08/23/2023 @ 11:15 |
| 08/23/2023 @ 11:45 | 09/18/2023 @ 11:45 |
| 09/18/2023 @ 12:00 | 10/30/2023 @ 09:45 |
| 10/30/2023 @ 10:15 | 11/28/2023 @ 10:30 |
| 11/28/2023 @ 10:45 | 12/29/2023 @ 10:15 |
| 12/29/2023 @ 10:45 | 01/09/2024 @ 11:30 |

Site: TN

| Deployment | Retrieval |
| --- | --- |
| NO INSTRUMENT DEPLOYED DUE TO ICE |  |
| 03/23/2023 @ 09:30 | 04/11/2023 @ 10:45 |
| 04/11/2023 @ 11:00 | 05/25/2023 @ 10:45 |
| 05/25/2023 @ 11:00 | 06/07/2023 @ 10:45 |
| 06/07/2023 @ 11:00 | 07/19/2023 @ 09:00 |
| 07/19/2023 @ 09:30 | 08/17/2023 @ 08:45 |
| 08/17/2023 @ 09:00 | 09/06/2023 @ 10:45 |
| 09/06/2023 @ 11:15 | 09/28/2023 @ 09:45 |
| 09/28/2023 @ 10:15 | 10/19/2023 @ 10:00 |
| 10/19/2023 @ 10:15 | 11/20/2023 @ 13:00 |
| 11/21/2023 @ 13:45 | 01/04/2024 @ 13:30 |

Site: TS

| Deployment | Retrieval |
| --- | --- |
| NO INSTRUMENT DEPLOYMENT DUE TO ICE |  |
| 03/22/2023 @ 10:00 | 04/12/2023 @ 11:45 |
| 04/12/2023 @ 12:00 | 05/25/2023 @ 10:15 |
| 05/25/2023 @ 10:30 | 06/07/2023 @ 10:15 |
| 06/07/2023 @ 10:30 | 07/19/2023 @ 09:45 |
| 07/19/2023 @ 10:00 | 08/17/2023 @ 09:15 |
| 08/17/2023 @ 09:30 | 09/06/2023 @ 10:30 |
| 09/06/2023 @ 10:45 | 10/19/2023 @ 10:30 |
| 10/19/2023 @ 10:45 | 11/20/2023 @ 13:15 |
| 11/20/2023 @ 13:30 | 01/05/2024 @ 11:45 |

Site: NP

| Deployment | Retrieval |
| --- | --- |
| 12/27/2022 @ 11:30 | 01/27/2023 @ 10:45 |
| 01/27/2023 @ 11:15 | 02/22/2023 @ 09:30 |
| 02/22/2023 @ 09:45 | 03/27/2023 @ 11:15 |
| 03/27/2023 @ 11:30 | 04/21/2023 @ 07:30 |
| 04/21/2023 @ 08:00 | 05/12/2023 @ 13:15 |
| 05/12/2023 @ 13:30 | 06/21/2023 @ 09:30 |
| 06/21/2023 @ 09:45 | 07/05/2023 @ 10:15 |
| 07/05/2023 @ 10:30 | 08/02/2023 @ 09:45 |
| 09/08/2023 @ 09:30 | 10/06/2023 @ 12:30 |
| 10/19/2023 @ 13:00 | 11/17/2023 @ 10:45 |
| 11/17/2023 @ 11:15 | 12/07/2023 @ 14:00 |
| 12/07/2023 @ 14:15 | 12/28/2023 @ 09:45 |
| 12/28/2023 @ 10:00 | 01/31/2024 @ 13:15 |

Site: BM

| Deployment | Retrieval |
| --- | --- |
| 12/20/2022 @ 13:30 | 01/19/2023 @ 12:00 |
| 01/19/2023 @ 12:15 | 02/14/2023 @ 11:30 |
| 02/14/2023 @ 11:45 | 03/28/2023 @ 10:45 |
| 03/28/2023 @ 11:00 | 04/27/2023 @ 11:00 |
| 04/27/2023 @ 11:15 | 05/31/2023 @ 12:00 |
| 05/31/2023 @ 12:30 | 06/29/2023 @ 10:00 |
| 06/29/2023 @ 10:15 | 08/10/2023 @ 11:30 |
| 08/10/2023 @ 11:45 | 09/26/2023 @ 11:15 |
| 09/26/2023 @ 11:45 | 10/23/2023 @ 11:00 |
| 10/23/2023 @ 11:15 | 11/21/2023 @ 11:45 |
| 11/21/2023 @ 12:00 | 12/20/2023 @ 11:00 |
| 12/20/2023 @ 11:30 | 01/22/2024 @ 12:30 |

7)  Distribution –

NOAA retains the right to analyze, synthesize and publish summaries of the NERRS System-wide Monitoring Program data.  The NERRS retains the right to be fully credited for having collected and processed the data.  Following academic courtesy standards, the NERR site where the data were collected should be contacted and fully acknowledged in any subsequent publications in which any part of the data are used.  The data set enclosed within this package/transmission is only as good as the quality assurance and quality control procedures outlined by the enclosed metadata reporting statement.  The user bears all responsibility for its subsequent use/misuse in any further analyses or comparisons.  The Federal government does not assume liability to the Recipient or third persons, nor will the Federal government reimburse or indemnify the Recipient for its liability due to any losses resulting in any way from the use of this data.

Requested citation format:

NOAA National Estuarine Research Reserve System (NERRS). System-wide Monitoring Program. Data accessed from the NOAA NERRS Centralized Data Management Office website: http://www.nerrsdata.org/; accessed 12 October 2022.

NERR water quality data and metadata can be obtained from the Research Coordinator at the individual NERR site (please see Principal Investigators and Contact Persons), from the Data Manager at the Centralized Data Management Office (please see personnel directory under the general information link on the CDMO home page) and online at the CDMO home page www.nerrsdata.org.  Data are available in comma delimited format.

8)  Associated researchers and projects –

As part of the SWMP long-term monitoring program, HUD NERR also collects 15-minute meteorological data and monthly grab and diel samples for nutrient/pigment data which may be correlated with this water quality dataset.  These data are available at www.nerrsdata.org.

II.  Physical Structure Descriptors

9)  Sensor specifications –

HUDNERR deployed YSI EXO 2 dataloggers at all sites (see Section 4).  All probes are titanium in construction and wiped by a central wiper unit.

Parameter: Temperature

Units: Celsius (C)

Sensor Type: Thermistor

Model#: 599870-01

Range: -5 to 50 C

Accuracy: -5 to 35: +/- 0.01, 35 to 50: +/- .005

Resolution: 0.01 C

Parameter: Conductivity

Units: milli-Siemens per cm (mS/cm)

Sensor Type: 4-electrode cell with autoranging

Model#: 599870-01

Range: 0 to 200 mS/cm

Accuracy: 0 to 100: +/- 0.5% of reading or 0.001 mS/cm; 100 to 200: +/- 1% of reading

Resolution: 0.001 mS/cm to 0.1 mS/cm (range dependent)Parameter: Salinity

Units: practical salinity units (psu)/parts per thousand (ppt)

Sensor Type: Calculated from conductivity and temperature

Range: 0 to 70 psu

Accuracy: +/- 1.0% of reading pr 0.1 ppt, whichever is greater

Resolution: 0.01 psu

WIPED CT SENSOR

Parameter: Temperature

Units: Celsius (C)

Sensor Type: Wiped probe; Thermistor

Model#: 599827

Range: -5 to 50 C

Accuracy: ±0.2 C

Resolution: 0.001 C

Parameter: Conductivity

Units: milli-Siemens per cm (mS/cm)

Sensor Type: Wiped probe; 4-electrode cell with autoranging

Model#: 599827

Range: 0 to 100 mS/cm

Accuracy: ±1% of the reading or 0.002 mS/cm, whichever is greater

Resolution: 0.0001 to 0.01 mS/cm (range dependent)

Parameter: Salinity

Units: practical salinity units (psu)/parts per thousand (ppt)

Model#: 599827

Sensor Type: Wiped probe; Calculated from conductivity and temperature

Range: 0 to 70 ppt

Accuracy: ±2% of the reading or 0.2 ppt, whichever is greater

Resolution: 0.01 psu

Parameter: Dissolved Oxygen % saturation

Sensor Type: Optical probe w/ mechanical cleaning

Model#: 599100-01

Range: 0 to 500% air saturation

Accuracy: 0-200% air saturation: +/- 1% of the reading or 1% air saturation, whichever is greater 200-500% air saturation: +/- 5% or reading

Resolution: 0.1% air saturation

Parameter: Dissolved Oxygen mg/L (Calculated from % air saturation, temperature, and salinity)

Units: milligrams/Liter (mg/L)

Sensor Type: Optical probe w/ mechanical cleaning

Model#: 599100-01

Range: 0 to 50 mg/L

Accuracy: 0-20 mg/L: +/-0.1 mg/l or 1% of the reading, whichever is greater

20 to 50 mg/L: +/- 5% of the reading

Resolution: 0.01 mg/L

Parameter: Non-vented Level - Shallow (Depth)

Units: feet or meters (ft or m)

Sensor Type: Stainless steel strain gauge

Range: 0 to 33 ft (10 m)

Accuracy: +/- 0.013 ft (0.004 m)

Resolution: 0.001 ft (0.001 m)

Parameter: pH

Units: pH units

Sensor Type: Glass combination electrode

Model#: 599701(guarded) or 599702(wiped)

Range: 0 to 14 units

Accuracy: +/- 0.01 units within +/- 10° of calibration temperature, +/- 0.02 units for entire temperature range

Resolution: 0.01 units

Parameter: Turbidity

Units: formazin nephelometric units (FNU)

Sensor Type: Optical, 90-degree scatter

Model#: 599101-01

Range: 0 to 4000 FNU

Accuracy: 0 to 999 FNU: 0.3 FNU or +/-2% of reading (whichever is greater); 1000 to 4000 FNU +/-5% of reading

Resolution: 0 to 999 FNU: 0.01 FNU, 1000 to 4000 FNU: 0.1 FNU

Parameter: Chlorophyll

Units: micrograms/Liter

Sensor Type: Optical probe

Model#: 599102-01

Range: 0 to 400 ug/Liter

Accuracy: Dependent on methodology

Resolution: 0.1 ug/L chl a, 0.1% FS

Depth Qualifier:

The NERR System-Wide Monitoring Program utilizes YSI data sondes that can be equipped with either vented or non-vented depth/level sensors.  Readings for both vented and non-vented sensors are automatically compensated for water density change due to variations in temperature and salinity; but for all non-vented depth measurements, changes in atmospheric pressure between calibrations appear as changes in water depth.  The error is equal to approximately 1.02 cm for every 1 millibar change in atmospheric pressure, and is eliminated for vented sensors because they are vented to the atmosphere throughout the deployment time interval.

Beginning in 2006, NERR SWMP standard calibration protocol calls for all non-vented depth sensors to read 0 meters at a (local) barometric pressure of 1013.25 mb (760 mm/hg).  To achieve this, each site calibrates their depth sensor with a depth offset number, which is calculated using the actual atmospheric pressure at the time of calibration and the equation provided in the SWMP calibration sheet or digital calibration log.  This offset procedure standardizes each depth calibration for the entire NERR System.  If accurate atmospheric pressure data are available, non-vented sensor depth measurements at any NERR can be corrected.

In 2010, the CDMO began automatically correcting Depth/Level data for changes in barometric pressure as measured by the reserve’s associated meteorological station during data ingestion.  These corrected Depth/Level data are reported as cDepth and cLevel, and are assigned QAQC flags and codes based on QAQC protocols.  Please see sections 11 and 12 for QAQC flag and code definitions.

NOTE:  older Depth data cannot be corrected without verifying that the depth offset was in place and whether a vented or non-vented depth sensor was in use.  No SWMP data prior to 2006 can be corrected using this method.  The following equation is used for corrected Depth/Level data provided by the CDMO beginning in 2010:

((1013-BP)*0.0102)+Depth/Level = cDepth/cLevel.

Salinity Units Qualifier:

In 2013, EXO sondes were approved for SWMP use and began to be utilized by reserves. While the 6600 series sondes report salinity in parts per thousand (ppt) units, the EXO sondes report practical salinity units (psu). These units are essentially the same and for SWMP purposes are understood to be equivalent, however psu is considered the more appropriate designation. Moving forward the NERR System will assign psu salinity units for all data regardless of sonde type.

Turbidity Qualifier:

In 2013, EXO sondes were approved for SWMP use and began to be utilized by reserves. While the 6600 series sondes report turbidity in nephelometric turbidity units (NTU), the EXO sondes use formazin nephelometric units (FNU). These units are essentially the same but indicate a difference in sensor methodology, for SWMP purposes they will be considered equivalent.  Moving forward, the NERR System will use FNU/NTU as the designated units for all turbidity data regardless of sonde type. If turbidity units and sensor methodology are of concern, please see the Sensor Specifications portion of the metadata.

Chlorophyll Fluorescence Disclaimer:

YSI chlorophyll sensors (6025 or 599102-01) are designed to serve as a proxy for chlorophyll concentrations in the field for monitoring applications and complement traditional lab extraction methods; therefore, there are accuracy limitations associated with the data that are detailed in the YSI manual including interference from other fluorescent species, differences in calibration method, and effects of cell structure, particle size, organism type, temperature, and light on sensor measurements.

10)  Coded variable definitions –

Sampling station:		Sampling site code:	Station code:

Ferry Landing		FL			hudflwq

Tivoli North Bay		TN			hudtnwq

Tivoli South Bay		TS			hudtswq

Norrie Point		NP			hudnpwq

Bear Mountain		BM			hudbmwq

11)  QAQC flag definitions –

QAQC flags provide documentation of the data and are applied to individual data points by insertion into the parameter’s associated flag column (header preceded by an F_).  During primary automated QAQC (performed by the CDMO), -5, -4, and -2 flags are applied automatically to indicate data that is missing and above or below sensor range.  All remaining data are then flagged 0, passing initial QAQC checks.   During secondary and tertiary QAQC 1, -3, and 5 flags may be used to note data as suspect, rejected due to QAQC, or corrected.

-5	Outside High Sensor Range

-4	Outside Low Sensor Range

-3		Data Rejected due to QAQC

-2		Missing Data

-1		Optional SWMP Supported Parameter

0		Data Passed Initial QAQC Checks

1		Suspect Data

2		Open - reserved for later flag

3		Calculated data: non-vented depth/level sensor correction for changes in barometric pressure

4		Historical Data:  Pre-Auto QAQC

5		Corrected Data

12)  QAQC code definitions –

QAQC codes are used in conjunction with QAQC flags to provide further documentation of the data and are also applied by insertion into the associated flag column.  There are three (3) different code categories, general, sensor, and comment.  General errors document general problems with the deployment or YSI datasonde, sensor errors are sensor specific, and comment codes are used to further document conditions or a problem with the data.  Only one general or sensor error and one comment code can be applied to a particular data point, but some comment codes (marked with an * below) can be applied to the entire record in the F_Record column.

General Errors

GIC	No instrument deployed due to ice

GIM	Instrument malfunction

GIT	Instrument recording error; recovered telemetry data

GMC 	No instrument deployed due to maintenance/calibration

GNF	Deployment tube clogged / no flow

GOW	Out of water event

GPF	Power failure / low battery

GQR	Data rejected due to QA/QC checks

GSM	See metadata

Corrected Depth/Level Data Codes

GCC	Calculated with data that were corrected during QA/QC

GCM	Calculated value could not be determined due to missing data

GCR	Calculated value could not be determined due to rejected data

GCS	Calculated value suspect due to questionable data

GCU 	Calculated value could not be determined due to unavailable data

Sensor Errors

SBO	Blocked optic

SCF	Conductivity sensor failure

SCS		Chlorophyll spike

SDF	Depth port frozen

SDG	Suspect due to sensor diagnostics

SDO	DO suspect

SDP	DO membrane puncture

SIC		Incorrect calibration / contaminated standard

SNV	Negative value

SOW	Sensor out of water

SPC	Post calibration out of range

SQR	Data rejected due to QAQC checks

SSD	Sensor drift

SSM	Sensor malfunction

SSR		Sensor removed / not deployed

STF	Catastrophic temperature sensor failure

STS		Turbidity spike

SWM	Wiper malfunction / loss

Comments

CAB*	Algal bloom

CAF	Acceptable calibration/accuracy error of sensor

CAP	Depth sensor in water, affected by atmospheric pressure

CBF	Biofouling

CCU	Cause unknown

CDA*	DO hypoxia (<3 mg/L)

CDB*	Disturbed bottom

CDF	Data appear to fit conditions

CFK*	Fish kill

CIP	*	Surface ice present at sample station

CLT*	Low tide

CMC*	In field maintenance/cleaning

CMD*	Mud in probe guard

CND	New deployment begins

CRE*	Significant rain event

CSM*	See metadata

CTS	Turbidity spike

CVT*	Possible vandalism/tampering

CWD*	Data collected at wrong depth

CWE*	Significant weather event

13)  Post deployment information –

FL

| Deploy Date | Sonde Nickname | SpCond | ROXDO1 | pH7 | pH10 | Turb | Turb | Depth |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 3/9/2023 | BALDUR | 1.083(1.0) | 100.3 | 7.17 | 10.09 | 1.02(0.0) | 112.65(124.0) | 0.121(0.122) |
| 4/20/2023 | LOKI | 0.951(1.0) | 101.6 | 7.08 | 10.06 | 4.24(0.0) | 123.71(124.0) | 0.184(0.15) |
| 5/19/2023 | FREYA | 0.938(1.0) | 101.6 | 7.08 | 9.92 | 2.98(0.0) | 118.73(124.0) | 0.073(0.122) |
| 6/20/2023 | KRATOS | 0.989(1.0) | 100.9 | 7.12 | 9.96 | 6.84(0.0) | 119.13(124.0) | 0.109(0.095) |
| 7/24/2023 | FREYA | 0.96(1.0) | 102 | 7.03 | 10.16 | 1.61(0.0) | 126.35(124.0) | 0.086(0.088) |
| 8/23/2023 | ODIN | 0.964(1.0) | 99.4 | 7.28 | 10.09 | 5.56(0.0) | 121.78(124.0) | -0.036(-0.014) |
| 9/18/2023 | KRATOS | 0.962(1.0) | 100.1 | 7.13 | 10.16 | 5.28(0.0) | 126.8(124.0) | 0.016(-0.027) |
| 10/30/2023 | BALDUR | 0.942(1.0) | 101.2 | 6.96 | 9.97 | 0.7(0.0) | 124.73(124.0) | -0.036(0.014) |

TN

| Deploy Date | Sonde Nickname | SpCond | ROXDO1 | pH7 | pH10 | Turb | Turb | Depth |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 3/23/2023 | FREYA | 0.926(1.0) | 101.8 | 7.09 | 10.07 | 1.67(0.0) | 120.11(124.0) | 0.014(0.014) |
| 4/11/2023 | JORMUNGANDR | 0.897(1.0) | 103.8 | 7.13 | 10.26 | 2.63(0.0) | 117.88(124.0) | 0.168(0.163) |
| 5/25/2023 | ODIN | 0.962(1.0) | 99.4 | 6.99 | 10.05 | 3.65(0.0) | 132.44(124.0) | -0.089(0.041) |
| 6/7/2023 | IDUNN | 0.889(1.0) | 101 | 7.28 | 10.03 | 7.08(0.0) | * | 0.145(0.095) |
| 7/19/2023 | LOKI | 0.945(1.0) | 100.5 | 7.04 | 9.86 | 2.31(0.0) | 137.78(124.0) | 0.034(-0.041) |
| 8/17/2023 | KRATOS | 0.963(1.0) | 101.6 | 7.06 | 10.28 | 2.79(0.0) | 125.3(124.0) | 0.018(0.054) |
| 9/6/2023 | ATREUS | 0.914(1.0) | ** | 7.15 | 10.09 | 1.98(0.0) | 122.51(124.0) | -0.204(0.163) |
| 9/28/2023 | FREYA | 0.968(1.0) | 100.3 | 7.85 | 10.55 | 118.72(0.0) | 118.72(124.0) | -0.45(0.095) |

TS

| Deploy Date | Sonde Nickname | SpCond | ROXDO1 | pH7 | pH10 | Turb | Turb | Depth |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 3/22/2023 | DURLIN | 0.918(1.0) | 100.1 | *** | *** | 1.29(0.0) | 124.58(124.0) | -0.048(-0.014) |
| 4/12/2023 | KRATOS | 0.941(1.0) | 102.1 | 7.15 | 10.1 | 2.83(0.0) | 117.46(124.0) | 0.181(0.163) |
| 5/23/2023 | LOKI | 0.975(1.0) | 98.7 | 6.93 | 10 | 4.01(0.0) | 143.21(124.0) | -0.097(0.0) |
| 6/7/2023 | JORMUNGANDR | 0.948(1.0) | 100.9 | 7.26 | 10.01 | 48.2(0.0) | 125.31(124.0) | -0.014(0.095) |
| 7/19/2023 | BALDUR | 0.915(1.0) | 99.2 | 7.06 | 9.91 | 1.15(0.0) | 127.47(124.0) | -0.028(-0.041) |
| 8/17/2023 | IDUNN | 0.863(1.0) | 100.2 | 7.01 | 10.24 | 1.74(0.0) | 125.56(124.0) | 0.013(0.041) |
| 9/6/2023 | MIMIR | 0.968(1.0) | 101.3 | 7.15 | 10.19 | 1.5(0.0) | 123.42(124.0) | -0.042(0.095) |
| 10/19/2023 | HEIMDALL | 0.921(1.0) | 102 | 7.05 | 10.06 | 1.65(0.0) | 119.11(124.0) | 0.203(0.231) |

NP

| Deploy Date | Sonde Nickname | SpCond | ROXDO1 | pH7 | pH10 | Turb | Turb | Depth |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1/27/2023 | LOKI | 0.982(1.0) | 99.9 | 7.15 | 10.04 | 0.22(0.0) | 113.92(124.0) | 0.044(0.041) |
| 2/22/2023 | IDUNN | 0.999(1.0) | 101.1 | 7.06 | 10.1 | -1.56(0.0) | 125.02(124.0) | -0.016(0.068) |
| 3/27/2023 | HEIMDALL | 0.945(1.0) | 101.4 | 7.26 | 10.2 | 3.8(0.0) | 118.49(124.0) | 0.093(0.095) |
| 4/21/2023 | ODIN | 0.946(1.0) | 100.1 | 7.09 | 10.13 | 0.96(0.0) | 120.83(124.0) | 0.081(0.088) |
| 5/12/2023 | DURLIN | 0.869(1.0) | 100.5 | 7.11 | 10.01 | 1.95(0.0) | 121.34(124.0) | 0.151(0.15) |
| 6/21/2023 | LOKI | 0.948(1.0) | 100.1 | 7.07 | 10.05 | 2.78(0.0) | 119.93(124.0) | 0.058(0.095) |
| 7/5/2023 | MIMIR | 0.917(1.0) | 101 | 7.09 | 10.07 | 2.06(0.0) | 123.38(124.0) | 0.167(0.082) |
| 8/2/2023 | HEIMDALL | 0.832(1.0) | 97.1 | 7.06 | 10.19 | 10.7(0.0) | 151.83(124.0) | -0.105(0.041) |

BM

| Deploy Date | Sonde Nickname | SpCond | ROXDO1 | pH7 | pH10 | Turb | Turb | Depth |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1/19/2023 | KRATOS | 10.177(10.0) | 101.9 | 7.08 | 10 | -0.03(0.0) | 122.26(124.0) | 0.13(0.068) |
| 2/14/2023 | ATREUS | *** | 102.5 | 7 | 10.01 | ** | ** | -0.021(0.082) |
| 3/28/2023 | HEL | 9.457(10.0) | 105.1 | 7.17 | 10.13 | 1.17(0.0) | -124 | -0.069(0.136) |
| 4/27/2023 | BALDUR | 9.505(10.0) | 100.3 | 7.09 | 10.08 | 0.87(0.0) | 116.49(124.0) | 0.094(0.109) |
| 5/31/2023 | HEIMDALL | 9.6(10.0) | 100.3 | 7.11 | 9.97 | 1.68(0.0) | 121.13(124.0) | 0.028(0.095) |
| 6/29/2023 | ODIN | 9.687(10.0) | 101 | 7.08 | 10.02 | 1.5(0.0) | 123.91(124.0) | 0.046(-0.054) |
| 8/10/2023 | JORMUNGANDR | 10.44(10.0) | 102.1 | 7.05 | 10.16 | 5.47(0.0) | 116.41(124.0) | 0.189(0.19) |
| 9/26/2023 | BALDUR | 9.411(10.0) | 101.9 | 7.1 | 10.14 | 3.36(0.0) | 124.56(124.0) | 0.205(0.19) |

* No value recorded.

** Sensor malfunctioning during post deployment checks. Data during the deployment do not appear to be affected. Cause of malfunction unknown.

*** No post deployment value due to sensor failure.

14)  Other remarks/notes –

Data are missing due to equipment or associated specific probes not being deployed, equipment failure, time of maintenance or calibration of equipment, or repair/replacement of a sampling station platform.  Any NANs in the dataset stand for “not a number” and are the result of low power, disconnected wires, or out of range readings.  If additional information on missing data is needed, contact the Research Coordinator at the reserve submitting the data.

Data Coded <CRE> or <CWE> Significant Rain or Significant Weather Event

Note: Due to the isolated nature of rain events, and geographic locations of sampling sites, HRNERR staff have determined that the FL and BM stations are not accurately reflected by either of our current meteorological monitoring stations. Except in extreme storm events (i.e., a hurricane), we will not be flagging CRE or CWE events until meteorological stations can be established nearby. An asterisk (*) in the table below indicates a CWE.

| NP (as recorded at NP MET) |  |
| --- | --- |
| 01/03/2023 @ 09:45 | 01/07/2023 @ 00:00 |
| 01/22/2023 @ 18:45 | 01/24/2023 @ 00:00 |
| 01/25/2023 @ 11:30 | 01/27/2023 @ 00:00 |
| 03/03/2023 @ 19:45 | 03/05/2023 @ 00:00 |
| 03/13/2023 @ 19:45 | 03/16/2023 @ 00:00* |
| 04/22/2023 @ 18:30 | 04/24/2023 @ 00:00 |
| 04/28/2023 @ 23:30 | 05/02/2023 @ 00:00 |
| 06/26/2023 @ 08:45 | 06/28/2023 @ 00:00 |
| 07/02/2023 @ 02:15 | 07/03/2023 @ 00:00 |
| 07/09/2023 @ 12:00 | 07/11/2023 @ 00:00 |
| 07/16/2023 @ 03:00 | 07/17/2023 @ 00:00 |
| 07/25/2023 @ 13:45 | 07/26/2023 @ 00:00 |
| 08/12/2023 @ 18:00 | 08/14/2023 @ 00:00 |
| 08/14/2023 @ 22:45 | 08/17/2023 @ 00:00 |
| 09/13/2023 @ 05:00 | 09/14/2023 @ 00:00 |
| 09/17/2023 @ 20:45 | 09/20/2023 @ 00:00 |
| 09/23/2023 @ 09:15 | 09/26/2023 @ 00:00 |
| 09/28/2023 @ 20:30 | 09/30/2023 @ 00:00 |
| 10/07/2023 @ 00:15 | 10/08/2023 @ 00:00 |
| 10/20/2023 @ 07:15 | 10/22/2023 @ 00:00 |
| 10/29/2023 @ 08:00 | 10/31/2023 @ 00:00 |
| 12/10/2023 @ 11:45 | 12/12/2023 @ 00:00 |
| 12/17/2023 @ 14:30 | 12/19/2023 @ 00:00 |

| TN (as recorded at Field Station (FS) MET) |  |
| --- | --- |
| 04/22/2023 @ 18:45 | 04/24/2023 @ 00:00 |
| 04/28/2023 @ 23:45 | 05/03/2023 @ 00:00 |
| 05/03/2023 @ 17:00 | 05/05/2023 @ 00:00 |
| 06/23/2023 @ 16:30 | 06/25/2023 @ 00:00 |
| 06/26/2023 @ 12:30 | 06/28/2023 @ 00:00 |
| 07/02/2023 @ 02:45 | 07/03/2023 @ 00:00 |
| 07/09/2023 @ 13:15 | 07/11/2023 @ 00:00 |
| 07/16/2023 @ 03:15 | 07/17/2023 @ 00:00 |
| 08/03/2023 @ 16:45 | 08/04/2023 @ 00:00 |
| 08/07/2023 @ 03:45 | 08/08/2023 @ 00:00 |
| 08/14/2023 @ 21:45 | 08/17/2023 @ 00:00 |
| 09/17/2023 @ 21:45 | 09/19/2023 @ 00:00 |
| 09/23/2023 @ 09:30 | 09/26/2023 @ 00:00 |
| 09/29/2023 @ 02:45 | 10/01/2023 @ 00:00 |
| 10/07/2023 @ 11:00 | 10/08/2023 @ 00:00 |
| 10/20/2023 @ 07:45 | 10/22/2023 @ 00:00 |
| 10/29/2023 @ 08:15 | 10/31/2023 @ 00:00 |
| 11/21/2023 @ 18:15 | 11/23/2023 @ 00:00 |
| 12/10/2023 @ 11:15 | 12/12/2023 @ 00:00 |
| 12/17/2023 @ 14:30 | 12/19/2023 @ 00:00 |

| TS (as recorded at FS MET) |  |
| --- | --- |
| 04/22/2023 @ 18:45 | 04/24/2023 @ 00:00 |
| 04/28/2023 @ 23:45 | 05/03/2023 @ 00:00 |
| 05/03/2023 @ 17:00 | 05/05/2023 @ 00:00 |
| 06/23/2023 @ 16:30 | 06/25/2023 @ 00:00 |
| 06/26/2023 @ 12:30 | 06/28/2023 @ 00:00 |
| 07/02/2023 @ 02:45 | 07/03/2023 @ 00:00 |
| 07/09/2023 @ 13:15 | 07/11/2023 @ 00:00 |
| 07/16/2023 @ 03:15 | 07/17/2023 @ 00:00 |
| 08/03/2023 @ 16:45 | 08/04/2023 @ 00:00 |
| 08/07/2023 @ 03:45 | 08/08/2023 @ 00:00 |
| 08/14/2023 @ 21:45 | 08/17/2023 @ 00:00 |
| 09/17/2023 @ 21:45 | 09/19/2023 @ 00:00 |
| 09/23/2023 @ 09:30 | 09/26/2023 @ 00:00 |
| 09/29/2023 @ 02:45 | 10/01/2023 @ 00:00 |
| 10/07/2023 @ 11:00 | 10/08/2023 @ 00:00 |
| 10/20/2023 @ 07:45 | 10/22/2023 @ 00:00 |
| 10/29/2023 @ 08:15 | 10/31/2023 @ 00:00 |
| 11/21/2023 @ 18:15 | 11/23/2023 @ 00:00 |
| 12/10/2023 @ 11:15 | 12/12/2023 @ 00:00 |
| 12/17/2023 @ 14:30 | 12/19/2023 @ 00:00 |

BM (07/09/2023 @ 00:00 – 07/30/2023 @ 00:00) (CRE)

All parameters during this timeframe were coded (CRE) due to the July 2023 Northeastern United States floods. Throughout this interval, there were multiple severe rain events (up to 266 mm) and flash flood events, which were localized in proximity to the BM station. Access to the site was inhibited due to catastrophic road destruction. Neither the station infrastructure nor the equipment experienced damages or interruptions in data acquisition. As there is no current NERRS meteorology monitoring station near this site, we are unable to identify exact rainfall amounts and durations; as such, the coded interval is based upon a generalized timeline of the event.

Data Coded (CWD) Data Collected at Wrong Depth

TN (09/28/2023 @ 10:15 – 10/02/2023 @ 10:00)

During the 9/28 deployment, every attempt was made to deploy the sonde properly in the sonde tube. Unfortunately, due to unusual high water related to the hurricane, the sonde was not situated within the deployment tube properly. Technicians returned to the station on 10/2 and found the sonde about 1 foot west of the deployment tube. They relocated it to the tube at that time, but during this interval data were not collected at the correct depth or according to NERR SWMP SOPs.

Data Coded [SQR] Data Rejected Due to QAQC Checks

TN (09/28/2023 @ 10:15 – 10/19/2023 @ 10:00)

Post-deployment checks are indicative of a pH sensor failure. Data appear to be disjointed from the previous and subsequent deployments. Additionally, data appear to be inconsistent with typical conditions of this site. As such, pH data have been rejected for the entire deployment.

Data Coded [SIC] Incorrect Calibration / Contaminated Standard

BM (11/21/2023 @ 12:00 – 12/20/2023 @ 11:00)

Chlorophyll data appear disjointed from the prior and subsequent deployments. No anomalies were identified upon review of the calibration data sheets. It is likely that the sensor was calibrated using contaminated standard, resulting in the offset data. As such, these data have been coded as suspect <1> [SIC].

NP (03/27/2023 @ 11:30 – 04/21/2023 @ 07:30)

pH data during this deployment have been rejected as values appear to be significantly higher than usual observations, and data are disjointed from the previous and subsequent deployments. A new pH module was installed onto the sensor, however, it appears that it may have been calibrated in a bad batch of standard

NP (09/08/2023 @ 09:30 – 10/06/2023 @ 12:30)

pH data have been flagged as suspect due to potential calibration in contaminated standard. Pre-deployment and post-deployment readings were slightly off, and the data appears disjointed from the previous and subsequent deployments.

NP (11/17/2023 @ 11:30 – 12/07/2023 @ 14:00)

DO and depth data have been flagged as suspect due to potential calibration error. It appears that the technician utilized an incorrect barometric pressure during the calibration process.

Data Coded [SSM] Sensor Malfunction

FL (06/20/2023 @ 11:00 – 07/24/2023 @ 10:45)

TN (08/17/2023 @ 09:00 – 09/06/2023 @ 10:45)

TS (04/12/2023 @ 12:15 - 05/25/2023 @ 10:15)

Conductivity and salinity data during these intervals indicate sensor dropout throughout the deployments. These are associated with a single wiped conductivity/temperature sensor (serial number 16H103200, model 599827). There appears to be a trend in the average values of the data which correspond with the prior and subsequent deployments, however, the sensor continuously “drops out”, resulting in atypical conductivity readings. Temperature readings appear to be unaffected. As such, temperature data were not flagged, and specific conductance and salinity are coded <1> [SSM]. This sensor has been retired and will not be utilized for SWMP use.

Data Coded [SDO] DO Suspect, [SSM] Sensor Malfunction

TN (09/10/2023 @ 02:15 – 09/28/2023 @ 09:45)

During this deployment, the DO sensor cut out intermittently. Data which are missing are coded as a sensor malfunction [SSM], while data which are present are coded as suspect DO <1>[SDO]. The cause of the sensor failure is likely due to age (2013 manufacture date).

Data Coded [SCF] Conductivity sensor Failure

TN (09/26/2023 @ 01:15 – 09/28/2023 @ 09:45)

During this interval there was a catastrophic failure of the conductivity/temperature probe. As a result, all dependent parameters (specific conductance, salinity, and depth) have been rejected. DO data are missing and are therefore coded <-2> (see above for data coded [SDO] during this interval). Although temperature data do not appear impacted, they have been coded <1> SSM due to periods of missing data during the deployment.

Data Coded [SWM] Wiper malfunction / Loss

NP (09/03/2023 @ 19:00 – 09/08/2023 @ 09:00)

Chlorophyll and turbidity data have been coded as suspect due to the loss of the central wiper brush during the deployment. The wiper brush was not recovered at the time of the instrument swap.

Data Coded [GIT] Instrument Recording Error; Recovered Telemetry Data

BM (01/01/2023 @ 00:00 – 01/19/2023 @ 12:00)

During the 12/20/2022 – 01/19/2023 deployment, the SD card on the instrument failed. Telemetered data was recovered in lieu of internally logged data, which was unrecoverable. Sporadic data points are missing, likely due to the reduced reliability of telemetered information. These data are flagged <-2> [GIT] (CSM), while recovered data are flagged <1> [GIT] (CSM).

TN (11/20/2023 @ 13:15 – 11/21/2023 @ 13:30)

Due to a deployment template error, the instrument did not internally record data for the duration of this deployment. Telemetered data were recovered for most of the deployment; however, data are missing from the interval 11/20/2023 @ 13:15 – 11/21/2023 @ 13:30.

Data Coded [GIM] Instrument Malfunction

BM (03/16/2023 @ 02:45 – 03/28/2023 @ 10:45)

Multiple sensors ceased to function correctly during this interval, resulting in missing data. Turbidity failed at 03/16/2023 @ 02:45, and the conductivity/temperature sensor failed at 03/18/2023 @ 00:45. All related parameters were rejected <-3> after 03/18/2023 @ 00:45. During post-deployment diagnostics, it was determined that port 3 on the sonde’s bulkhead had failed, resulting in the malfunctioning sensors. Data values were recorded until 03/20/2023 @ 02:30, after which the instrument malfunctioned, and no data were reported.

Data Coded [SPC] Post Calibration Out of Range

TN (09/28/2023 @ 10:15 – 10/19/2023 @ 10:00)

pH data appear disjointed from the prior and subsequent deployments. Post-deployment checks for the sensor were elevated beyond typical readings. pH for the entire deployment has been coded as suspect <1> [SPC].

https://extapps.dec.ny.gov/cfmx/extapps/derexternal/haz/details.cfm?ProgNo=546031

https://www.epa.gov/system/files/documents/2023-05/English.pdf

https://streamstats.usgs.gov/ss/

Wikipedia contributors. (2023, October 9). July 2023 Northeastern United States floods. In Wikipedia, The Free Encyclopedia. Retrieved 14:06, October 19, 2023, from https://en.wikipedia.org/w/index.php?title=July_2023_Northeastern_United_States_floods&oldid=1179352801